## Sertifikat
SERTIFIKAT TABEL VOLUME TANGKI
Nomor: 3816/PKTN.4.8/KHP/DL/10/2019
[7 empty] | NOMOR TANGKI |  | : | T00902
[7 empty] | JENIS ATAP |  | : | Tetap
[7 empty] | DIAMETER |  | : | 34820 | mm
[7 empty] | VOLUME BERSIH |  | : | 10783279 | liter
[7 empty] | PEMILIK |  | : | PT. Pertamina (Persero) RU V
[7 empty] | LOKASI |  | : | Balikpapan - Kalimantan Timur
ELEVASI :
[7 empty] | Tinggi lubang ukur dari meja ukur |  |  |  | AD = | 12671 | mm
[7 empty] | Tinggi lubang ukur dari dasar |  |  |  | AE = | 12753 | mm
[7 empty] | Tinggi tangki |  |  |  | BE = | 11818 | mm
[7 empty] | Tinggi maksimum volume bersih |  |  |  | CE = | 11500 | mm
[7 empty] | Tinggi meja ukur |  |  |  | DE = | 82 | mm
[7 empty] | Tinggi dasar tangki |  |  |  | E = | 0 | mm
| CATATAN : | [5 empty] | Ditera ulang :
[7 empty] | Tanggal : |  | 2019-10-11
| 1. | Tabel volume tangki ini dibuat untuk |  |  |  |  | Oleh : |  | 1. Hikmat Rijadi
|  | suhu |  | 36 | °C |  |  |  | 2. Anditya Kristianto
massa jenis cairan pada
|  | 36 °C = |  |  | 0.86 | g/ml
| 2. | Selain suhu 36 °C penunjukan tabel |  |  |  |  | DISAHKAN BERDASARKAN UNDANG UNDANG RI NOMOR 2 TAHUN 1981 TENTANG METROLOGI LEGAL DENGAN MEMBUBUHKAN TANDA TERA SAH PADA LEMPING VOLUME NOMINAL.
volume tangki harus dikalikan faktor:
|  | { 1 + |  | a | ( t - 36°C ) }
|  | t = | suhu tangki
|  | a = | koefisien muai ruang bahan
|  |  | dinding tangki per °C |  |  |  |  | Bandung, 29 Oktober 2019
| 3. | Tabel volume tangki ini berlaku | [5 empty] | KEMENTERIAN PERDAGANGAN R.I.
|  | 10 (sepuluh) tahun | [5 empty] | Direktorat Jenderal Perlindungan Konsumen dan Tertib Niaga
[7 empty] | a.n | Direktur Metrologi
| 4. | Tangki ukur ini agar ditera ulang | [5 empty] | Kepala Balai Pengujian UTTP
|  | paling lambat bulan: |  |  |  | Oktober 2029
| 5. | Penyerahan/penerimaan isi tangki
|  | diizinkan minimum sampai perubahan | [5 empty] | Eko Agus Irianto
|  | tinggi cairan 2 meter. | [5 empty] | NIP. 196208011984011002
NIP. 19661129 199603 1 001
K17-07-01-02 | [10 empty] | Nomor Order : |  | DL-352-17
A
B
C
D
E

## Hal Contoh 1
Halaman - 2 |  |  |  | Tabel Volume Tangki Nomor : 3816/PKTN.4.8/KHP/DL/10/2019
TETAP
CONTOH PEMAKAIAN TABEL VOLUME TANGKI
Tangki No.T-902
Umpama menghitung volume tangki No. T-902 dalam transaksi Penyerahan
Suhu dinding tangki selama Penyerahan rata-rata 35 °C
Tinggi cairan dari meja ukur menurut alat ukur tinggi/dip tape
| sebelum cairan diserahkan (h1) | = | 6007 | mm
| Maka tinggi cairan dari dasar tangki (h1+meja ukur) | = | 6089 | mm
| Volume cairan pada suhu 36 °C dibaca pada tabel volume tangki | = | 5636301 | liter
( 5627748 liter halaman 6 + 8553.2 liter cincin 5 halaman 3 )
Tinggi cairan dari meja ukur menurut alat ukur tinggi/dip tape
| setelah cairan diserahkan (h2) | = | 3003 | mm
| Maka tinggi cairan dari dasar tangki (h2+meja ukur) | = | 3085 | mm
| Volume cairan pada suhu 36 °C dibaca pada tabel volume tangki | = | 2780803 | liter
( 2776045 liter halaman 5 + 4757.7 liter cincin 3 halaman 3 )
Volume cairan yang diserahkan pada suhu 36 °C :
| 5636301 liter - 2780803 liter | = | 2855498 | liter
Faktor koreksi volume tangki akibat perubahan suhu
| dari 36 °C menjadi 35 °C adalah : 1 + { 0,0000348 x ( 35 - 36 ) } | = | 0.99996
Jadi volume cairan yang diserahkan pada suhu 35 °C
| 2855498 liter x 0.999965 | = | 2855399 | liter

## Fraksi
Halaman - 3 | [20 empty] | Tabel Volume Tangki Nomor : 3816/PKTN.4.8/KHP/DL/10/2019
TABEL FRAKSI DARI 1 mm SAMPAI 10 mm (Liter)
Tangki No.T-902
| Cincin | : | 1 |  | Cincin | : | 2 |  | Cincin | : | 3 |  | Cincin | : | 4 |  | Cincin | : | 4
| dari | : | 0.5 | m | dari | : | 1.68 | m | dari | : | 2.91 | m | dari | : | 4.45 | m | dari | : | 5.66 | m
| sampai | : | 1.68 | m | sampai | : | 2.91 | m | sampai | : | 4.45 | m | sampai | : | 5.66 | m | sampai | : | 7.51 | m
| mm |  | liter |  | mm |  | liter |  | mm |  | liter |  | mm |  | liter |  | mm |  | liter
| 1 |  | 951.3 |  | 1 |  | 949.3 |  | 1 |  | 951.5 |  | 1 |  | 949.5 |  | 1 |  | 950.4
| 2 |  | 1902.6 |  | 2 |  | 1898.6 |  | 2 |  | 1903.1 |  | 2 |  | 1899.1 |  | 2 |  | 1900.7
| 3 |  | 2854 |  | 3 |  | 2847.9 |  | 3 |  | 2854.6 |  | 3 |  | 2848.6 |  | 3 |  | 2851.1
| 4 |  | 3805.3 |  | 4 |  | 3797.2 |  | 4 |  | 3806.2 |  | 4 |  | 3798.1 |  | 4 |  | 3801.4
| 5 |  | 4756.6 |  | 5 |  | 4746.5 |  | 5 |  | 4757.7 |  | 5 |  | 4747.7 |  | 5 |  | 4751.8
| 6 |  | 5707.9 |  | 6 |  | 5695.8 |  | 6 |  | 5709.3 |  | 6 |  | 5697.2 |  | 6 |  | 5702.1
| 7 |  | 6659.2 |  | 7 |  | 6645.1 |  | 7 |  | 6660.8 |  | 7 |  | 6646.7 |  | 7 |  | 6652.5
| 8 |  | 7610.6 |  | 8 |  | 7594.4 |  | 8 |  | 7612.4 |  | 8 |  | 7596.3 |  | 8 |  | 7602.9
| 9 |  | 8561.9 |  | 9 |  | 8543.7 |  | 9 |  | 8563.9 |  | 9 |  | 8545.8 |  | 9 |  | 8553.2
| 10 |  | 9513.2 |  | 10 |  | 9493 |  | 10 |  | 9515.5 |  | 10 |  | 9495.3 |  | 10 |  | 9503.6
| Cincin | : | 6 |  | Cincin | : | 7 |  | Cincin | : | 8 |  | Cincin | : | - |  | Cincin | : | -
| dari | : | 7.51 | m | dari | : | 9.34 | m | dari | : | 11.18 | m | dari | : | - | m | dari | : | - | m
| sampai | : | 9.34 | m | sampai | : | 11.18 | m | sampai | : | 11.5 | m | sampai | : | - | m | sampai | : | - | m
| mm |  | liter |  | mm |  | liter |  | mm |  | liter |  | mm |  | liter |  | mm |  | liter
| 1 |  | 950.8 |  | 1 |  | 952 |  | 1 |  | 952.7
| 2 |  | 1901.7 |  | 2 |  | 1904 |  | 2 |  | 1903.3
| 3 |  | 2852.5 |  | 3 |  | 2856 |  | 3 |  | 2858
| 4 |  | 3803.3 |  | 4 |  | 3808 |  | 4 |  | 3810.6
| 5 |  | 4754.1 |  | 5 |  | 4759.9 |  | 5 |  | 4763.3
| 6 |  | 5705 |  | 6 |  | 5711.9 |  | 6 |  | 5715.9
| 7 |  | 6655.8 |  | 7 |  | 6663.9 |  | 7 |  | 6668.6
| 8 |  | 7606.6 |  | 8 |  | 7615.9 |  | 8 |  | 7621.2
| 9 |  | 8557.4 |  | 9 |  | 8567.9 |  | 9 |  | 8573.9
| 10 |  | 9508.3 |  | 10 |  | 9519.9 |  | 10 |  | 9526.5
Note on A17: VR'46: Halaman yang akan di print dipilih antara page 1, page 2 atau page 3 (disesuaikan dengan jumlah cincin/kolom yang terisi)

## Tabel
Halaman - |  |  | 4 | Nomor Tangki : T-902 | [6 empty] | Tabel Volume Tangki Nomor : 3816/PKTN.4.8/KHP/DL/10/2019 | [8 empty] | 700 | Halaman - |  |  | 5 | Nomor Tangki : T-902 | [6 empty] | 0 | [9 empty] | Halaman - |  |  | 6 | Nomor Tangki : T-902 | [6 empty] | 0
| Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume |  | Tinggi |  | Volume
m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter |  | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter |  | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter |  | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter | m |  | cm | liter
| 1 |  |  | 0 | 51 | 51 | 333773 | 1 | 101 | 1 | 809468 | 1 | 151 | 51 | 1284976 | 2 | 201 | 1 | 1759916 | 700 | 2 | 251 | 51 | 2234564 | 3 | 301 | 1 | 2709437 | 3 | 351 | 51 | 3185209 | 4 | 401 | 1 | 3660982 | 4 | 451 | 51 | 4136634 | 700 | 5 | 501 | 1 | 4611401 | 5 | 551 | 51 | 5086168 | 6 | 601 | 1 | 5561223 | 6 | 651 | 51 | 6036401 | 7 | 701 | 1 | 6511580 | 1340 | 7 | 751 | 51 | 6986758 | 8 | 801 | 1 | 7462171 | 8 | 851 | 51 | 7937585 | 9 | 901 | 1 | 8412998 | 9 | 951 | 51 | 8888609 | 10 | 1001 | 1 | 9364604 | 10 | 1051 | 51 | 9840598 | 11 | 1101 | 1 | 10316593 | 1151 | 1201 | 2150
| 2 |  |  | 0 | 52 | 52 | 343286 | 1 | 102 | 2 | 818978 | 1 | 152 | 52 | 1294486 | 2 | 202 | 2 | 1769409 | 700 | 2 | 252 | 52 | 2244057 | 3 | 302 | 2 | 2718952 | 3 | 352 | 52 | 3194725 | 4 | 402 | 2 | 3670497 | 4 | 452 | 52 | 4146129 | 700 | 5 | 502 | 2 | 4620896 | 5 | 552 | 52 | 5095663 | 6 | 602 | 2 | 5570726 | 6 | 652 | 52 | 6045905 | 7 | 702 | 2 | 6521083 | 1340 | 7 | 752 | 52 | 6996266 | 8 | 802 | 2 | 7471680 | 8 | 852 | 52 | 7947093 | 9 | 902 | 2 | 8422506 | 9 | 952 | 52 | 8898129 | 10 | 1002 | 2 | 9374124 | 10 | 1052 | 52 | 9850118 | 11 | 1102 | 2 | 10326113 | 1152 | 1202 | 2150
| 3 |  |  | 0 | 53 | 53 | 352800 | 1 | 103 | 3 | 828489 | 1 | 153 | 53 | 1303996 | 2 | 203 | 3 | 1778902 | 700 | 2 | 253 | 53 | 2253550 | 3 | 303 | 3 | 2728467 | 3 | 353 | 53 | 3204240 | 4 | 403 | 3 | 3680013 | 4 | 453 | 53 | 4155625 | 700 | 5 | 503 | 3 | 4630392 | 5 | 553 | 53 | 5105159 | 6 | 603 | 3 | 5580230 | 6 | 653 | 53 | 6055408 | 7 | 703 | 3 | 6530587 | 1340 | 7 | 753 | 53 | 7005775 | 8 | 803 | 3 | 7481188 | 8 | 853 | 53 | 7956601 | 9 | 903 | 3 | 8432015 | 9 | 953 | 53 | 8907649 | 10 | 1003 | 3 | 9383644 | 10 | 1053 | 53 | 9859638 | 11 | 1103 | 3 | 10335633 | 1153 | 1203 | 2150
| 4 |  |  | 0 | 54 | 54 | 362313 | 1 | 104 | 4 | 837999 | 1 | 154 | 54 | 1313506 | 2 | 204 | 4 | 1788395 | 700 | 2 | 254 | 54 | 2263043 | 3 | 304 | 4 | 2737983 | 3 | 354 | 54 | 3213756 | 4 | 404 | 4 | 3689528 | 4 | 454 | 54 | 4165120 | 700 | 5 | 504 | 4 | 4639887 | 5 | 554 | 54 | 5114654 | 6 | 604 | 4 | 5589734 | 6 | 654 | 54 | 6064912 | 7 | 704 | 4 | 6540090 | 1340 | 7 | 754 | 54 | 7015283 | 8 | 804 | 4 | 7490696 | 8 | 854 | 54 | 7966110 | 9 | 904 | 4 | 8441523 | 9 | 954 | 54 | 8917169 | 10 | 1004 | 4 | 9393163 | 10 | 1054 | 54 | 9869158 | 11 | 1104 | 4 | 10345153 | 1154 | 1204 | 2150
| 5 |  |  | 0 | 55 | 55 | 371827 | 1 | 105 | 5 | 847509 | 1 | 155 | 55 | 1323016 | 2 | 205 | 5 | 1797887 | 700 | 2 | 255 | 55 | 2272535 | 3 | 305 | 5 | 2747498 | 3 | 355 | 55 | 3223271 | 4 | 405 | 5 | 3699044 | 4 | 455 | 55 | 4174615 | 700 | 5 | 505 | 5 | 4649382 | 5 | 555 | 55 | 5124149 | 6 | 605 | 5 | 5599237 | 6 | 655 | 55 | 6074416 | 7 | 705 | 5 | 6549594 | 1340 | 7 | 755 | 55 | 7024791 | 8 | 805 | 5 | 7500204 | 8 | 855 | 55 | 7975618 | 9 | 905 | 5 | 8451031 | 9 | 955 | 55 | 8926689 | 10 | 1005 | 5 | 9402683 | 10 | 1055 | 55 | 9878678 | 11 | 1105 | 5 | 10354672 | 1155 | 1205 | 2150
| 6 |  |  | 0 | 56 | 56 | 381341 | 1 | 106 | 6 | 857019 | 1 | 156 | 56 | 1332526 | 2 | 206 | 6 | 1807380 | 700 | 2 | 256 | 56 | 2282028 | 3 | 306 | 6 | 2757014 | 3 | 356 | 56 | 3232787 | 4 | 406 | 6 | 3708559 | 4 | 456 | 56 | 4184111 | 700 | 5 | 506 | 6 | 4658878 | 5 | 556 | 56 | 5133645 | 6 | 606 | 6 | 5608741 | 6 | 656 | 56 | 6083919 | 7 | 706 | 6 | 6559097 | 1340 | 7 | 756 | 56 | 7034299 | 8 | 806 | 6 | 7509713 | 8 | 856 | 56 | 7985126 | 9 | 906 | 6 | 8460540 | 9 | 956 | 56 | 8936209 | 10 | 1006 | 6 | 9412203 | 10 | 1056 | 56 | 9888198 | 11 | 1106 | 6 | 10364192 | 1156 | 1206 | 2150
| 7 |  |  | 0 | 57 | 57 | 390854 | 1 | 107 | 7 | 866529 | 1 | 157 | 57 | 1342036 | 2 | 207 | 7 | 1816873 | 700 | 2 | 257 | 57 | 2291521 | 3 | 307 | 7 | 2766529 | 3 | 357 | 57 | 3242302 | 4 | 407 | 7 | 3718075 | 4 | 457 | 57 | 4193606 | 700 | 5 | 507 | 7 | 4668373 | 5 | 557 | 57 | 5143140 | 6 | 607 | 7 | 5618244 | 6 | 657 | 57 | 6093423 | 7 | 707 | 7 | 6568601 | 1340 | 7 | 757 | 57 | 7043808 | 8 | 807 | 7 | 7519221 | 8 | 857 | 57 | 7994634 | 9 | 907 | 7 | 8470048 | 9 | 957 | 57 | 8945729 | 10 | 1007 | 7 | 9421723 | 10 | 1057 | 57 | 9897718 | 11 | 1107 | 7 | 10373712 | 1157 | 1207 | 2150
| 8 |  |  | 0 | 58 | 58 | 400368 | 1 | 108 | 8 | 876039 | 1 | 158 | 58 | 1351547 | 2 | 208 | 8 | 1826366 | 700 | 2 | 258 | 58 | 2301014 | 3 | 308 | 8 | 2776045 | 3 | 358 | 58 | 3251817 | 4 | 408 | 8 | 3727590 | 4 | 458 | 58 | 4203101 | 700 | 5 | 508 | 8 | 4677868 | 5 | 558 | 58 | 5152635 | 6 | 608 | 8 | 5627748 | 6 | 658 | 58 | 6102926 | 7 | 708 | 8 | 6578105 | 1340 | 7 | 758 | 58 | 7053316 | 8 | 808 | 8 | 7528729 | 8 | 858 | 58 | 8004143 | 9 | 908 | 8 | 8479556 | 9 | 958 | 58 | 8955248 | 10 | 1008 | 8 | 9431243 | 10 | 1058 | 58 | 9907238 | 11 | 1108 | 8 | 10383232 | 1158 | 1208 | 2150
| 9 |  |  | 0 | 59 | 59 | 409882 | 1 | 109 | 9 | 885549 | 1 | 159 | 59 | 1361057 | 2 | 209 | 9 | 1835859 | 700 | 2 | 259 | 59 | 2310507 | 3 | 309 | 9 | 2785560 | 3 | 359 | 59 | 3261333 | 4 | 409 | 9 | 3737106 | 4 | 459 | 59 | 4212597 | 700 | 5 | 509 | 9 | 4687364 | 5 | 559 | 59 | 5162131 | 6 | 609 | 9 | 5637251 | 6 | 659 | 59 | 6112430 | 7 | 709 | 9 | 6587608 | 1340 | 7 | 759 | 59 | 7062824 | 8 | 809 | 9 | 7538238 | 8 | 859 | 59 | 8013651 | 9 | 909 | 9 | 8489064 | 9 | 959 | 59 | 8964768 | 10 | 1009 | 9 | 9440763 | 10 | 1059 | 59 | 9916757 | 11 | 1109 | 9 | 10392752 | 1159 | 1209 | 2150
| 10 |  |  | 0 | 60 | 60 | 419395 | 1 | 110 | 10 | 895060 | 1 | 160 | 60 | 1370567 | 2 | 210 | 10 | 1845352 | 700 | 2 | 260 | 60 | 2320000 | 3 | 310 | 10 | 2795076 | 3 | 360 | 60 | 3270848 | 4 | 410 | 10 | 3746621 | 4 | 460 | 60 | 4222092 | 700 | 5 | 510 | 10 | 4696859 | 5 | 560 | 60 | 5171626 | 6 | 610 | 10 | 5646755 | 6 | 660 | 60 | 6121933 | 7 | 710 | 10 | 6597112 | 1340 | 7 | 760 | 60 | 7072332 | 8 | 810 | 10 | 7547746 | 8 | 860 | 60 | 8023159 | 9 | 910 | 10 | 8498573 | 9 | 960 | 60 | 8974288 | 10 | 1010 | 10 | 9450283 | 10 | 1060 | 60 | 9926277 | 11 | 1110 | 10 | 10402272 | 1160 | 1210 | 2150
| 11 |  |  | 0 | 61 | 61 | 428909 | 1 | 111 | 11 | 904570 | 1 | 161 | 61 | 1380077 | 2 | 211 | 11 | 1854845 | 700 | 2 | 261 | 61 | 2329493 | 3 | 311 | 11 | 2804591 | 3 | 361 | 61 | 3280364 | 4 | 411 | 11 | 3756137 | 4 | 461 | 61 | 4231587 | 700 | 5 | 511 | 11 | 4706354 | 5 | 561 | 61 | 5181121 | 6 | 611 | 11 | 5656259 | 6 | 661 | 61 | 6131437 | 7 | 711 | 11 | 6606615 | 1340 | 7 | 761 | 61 | 7081841 | 8 | 811 | 11 | 7557254 | 8 | 861 | 61 | 8032667 | 9 | 911 | 11 | 8508081 | 9 | 961 | 61 | 8983808 | 10 | 1011 | 11 | 9459803 | 10 | 1061 | 61 | 9935797 | 11 | 1111 | 11 | 10411792 | 1161 | 1211 | 2150
| 12 |  |  | 0 | 62 | 62 | 438422 | 1 | 112 | 12 | 914080 | 1 | 162 | 62 | 1389587 | 2 | 212 | 12 | 1864338 | 700 | 2 | 262 | 62 | 2338986 | 3 | 312 | 12 | 2814107 | 3 | 362 | 62 | 3289879 | 4 | 412 | 12 | 3765652 | 4 | 462 | 62 | 4241083 | 700 | 5 | 512 | 12 | 4715850 | 5 | 562 | 62 | 5190617 | 6 | 612 | 12 | 5665762 | 6 | 662 | 62 | 6140940 | 7 | 712 | 12 | 6616119 | 1340 | 7 | 762 | 62 | 7091349 | 8 | 812 | 12 | 7566762 | 8 | 862 | 62 | 8042176 | 9 | 912 | 12 | 8517589 | 9 | 962 | 62 | 8993328 | 10 | 1012 | 12 | 9469323 | 10 | 1062 | 62 | 9945317 | 11 | 1112 | 12 | 10421312 | 1162 | 1212 | 2150
| 13 |  |  | 0 | 63 | 63 | 447936 | 1 | 113 | 13 | 923590 | 1 | 163 | 63 | 1399097 | 2 | 213 | 13 | 1873831 | 700 | 2 | 263 | 63 | 2348479 | 3 | 313 | 13 | 2823622 | 3 | 363 | 63 | 3299395 | 4 | 413 | 13 | 3775167 | 4 | 463 | 63 | 4250578 | 700 | 5 | 513 | 13 | 4725345 | 5 | 563 | 63 | 5200112 | 6 | 613 | 13 | 5675266 | 6 | 663 | 63 | 6150444 | 7 | 713 | 13 | 6625622 | 1340 | 7 | 763 | 63 | 7100857 | 8 | 813 | 13 | 7576271 | 8 | 863 | 63 | 8051684 | 9 | 913 | 13 | 8527097 | 9 | 963 | 63 | 9002848 | 10 | 1013 | 13 | 9478842 | 10 | 1063 | 63 | 9954837 | 11 | 1113 | 13 | 10430832 | 1163 | 1213 | 2150
| 14 |  |  | 0 | 64 | 64 | 457450 | 1 | 114 | 14 | 933100 | 1 | 164 | 64 | 1408607 | 2 | 214 | 14 | 1883324 | 700 | 2 | 264 | 64 | 2357972 | 3 | 314 | 14 | 2833137 | 3 | 364 | 64 | 3308910 | 4 | 414 | 14 | 3784683 | 4 | 464 | 64 | 4260073 | 700 | 5 | 514 | 14 | 4734840 | 5 | 564 | 64 | 5209607 | 6 | 614 | 14 | 5684769 | 6 | 664 | 64 | 6159948 | 7 | 714 | 14 | 6635126 | 1340 | 7 | 764 | 64 | 7110365 | 8 | 814 | 14 | 7585779 | 8 | 864 | 64 | 8061192 | 9 | 914 | 14 | 8536606 | 9 | 964 | 64 | 9012368 | 10 | 1014 | 14 | 9488362 | 10 | 1064 | 64 | 9964357 | 11 | 1114 | 14 | 10440351 | 1164 |  | 2150
| 15 |  |  | 0 | 65 | 65 | 466963 | 1 | 115 | 15 | 942610 | 1 | 165 | 65 | 1418118 | 2 | 215 | 15 | 1892817 | 700 | 2 | 265 | 65 | 2367465 | 3 | 315 | 15 | 2842653 | 3 | 365 | 65 | 3318426 | 4 | 415 | 15 | 3794198 | 4 | 465 | 65 | 4269569 | 700 | 5 | 515 | 15 | 4744336 | 5 | 565 | 65 | 5219103 | 6 | 615 | 15 | 5694273 | 6 | 665 | 65 | 6169451 | 7 | 715 | 15 | 6644630 | 1340 | 7 | 765 | 65 | 7119874 | 8 | 815 | 15 | 7595287 | 8 | 865 | 65 | 8070701 | 9 | 915 | 15 | 8546114 | 9 | 965 | 65 | 9021888 | 10 | 1015 | 15 | 9497882 | 10 | 1065 | 65 | 9973877 | 11 | 1115 | 15 | 10449871 | 1165 |  | 2150
| 16 |  |  | 0 | 66 | 66 | 476477 | 1 | 116 | 16 | 952120 | 1 | 166 | 66 | 1427628 | 2 | 216 | 16 | 1902310 | 700 | 2 | 266 | 66 | 2376958 | 3 | 316 | 16 | 2852168 | 3 | 366 | 66 | 3327941 | 4 | 416 | 16 | 3803714 | 4 | 466 | 66 | 4279064 | 700 | 5 | 516 | 16 | 4753831 | 5 | 566 | 66 | 5228598 | 6 | 616 | 16 | 5703776 | 6 | 666 | 66 | 6178955 | 7 | 716 | 16 | 6654133 | 1340 | 7 | 766 | 66 | 7129382 | 8 | 816 | 16 | 7604795 | 8 | 866 | 66 | 8080209 | 9 | 916 | 16 | 8555622 | 9 | 966 | 66 | 9031408 | 10 | 1016 | 16 | 9507402 | 10 | 1066 | 66 | 9983397 | 11 | 1116 | 16 | 10459391 | 1166 |  | 2150
| 17 |  |  | 0 | 67 | 67 | 485991 | 1 | 117 | 17 | 961631 | 1 | 167 | 67 | 1437138 | 2 | 217 | 17 | 1911803 | 700 | 2 | 267 | 67 | 2386451 | 3 | 317 | 17 | 2861684 | 3 | 367 | 67 | 3337457 | 4 | 417 | 17 | 3813229 | 4 | 467 | 67 | 4288559 | 700 | 5 | 517 | 17 | 4763326 | 5 | 567 | 67 | 5238102 | 6 | 617 | 17 | 5713280 | 6 | 667 | 67 | 6188458 | 7 | 717 | 17 | 6663637 | 1340 | 7 | 767 | 67 | 7138890 | 8 | 817 | 17 | 7614304 | 8 | 867 | 67 | 8089717 | 9 | 917 | 17 | 8565130 | 9 | 967 | 67 | 9040927 | 10 | 1017 | 17 | 9516922 | 10 | 1067 | 67 | 9992917 | 11 | 1117 | 17 | 10468911 | 1167 |  | 2150
| 18 |  |  | 0 | 68 | 68 | 495504 | 1 | 118 | 18 | 971141 | 1 | 168 | 68 | 1446648 | 2 | 218 | 18 | 1921296 | 700 | 2 | 268 | 68 | 2395944 | 3 | 318 | 18 | 2871199 | 3 | 368 | 68 | 3346972 | 4 | 418 | 18 | 3822745 | 4 | 468 | 68 | 4298055 | 700 | 5 | 518 | 18 | 4772822 | 5 | 568 | 68 | 5247605 | 6 | 618 | 18 | 5722784 | 6 | 668 | 68 | 6197962 | 7 | 718 | 18 | 6673140 | 1340 | 7 | 768 | 68 | 7148399 | 8 | 818 | 18 | 7623812 | 8 | 868 | 68 | 8099225 | 9 | 918 | 18 | 8574639 | 9 | 968 | 68 | 9050447 | 10 | 1018 | 18 | 9526442 | 10 | 1068 | 68 | 10002436 | 11 | 1118 | 18 | 10478431 | 1168 |  | 2150
| 19 |  |  | 0 | 69 | 69 | 505018 | 1 | 119 | 19 | 980651 | 1 | 169 | 69 | 1456141 | 2 | 219 | 19 | 1930789 | 700 | 2 | 269 | 69 | 2405437 | 3 | 319 | 19 | 2880715 | 3 | 369 | 69 | 3356487 | 4 | 419 | 19 | 3832260 | 4 | 469 | 69 | 4307550 | 700 | 5 | 519 | 19 | 4782317 | 5 | 569 | 69 | 5257109 | 6 | 619 | 19 | 5732287 | 6 | 669 | 69 | 6207465 | 7 | 719 | 19 | 6682644 | 1340 | 7 | 769 | 69 | 7157907 | 8 | 819 | 19 | 7633320 | 8 | 869 | 69 | 8108734 | 9 | 919 | 19 | 8584147 | 9 | 969 | 69 | 9059967 | 10 | 1019 | 19 | 9535962 | 10 | 1069 | 69 | 10011956 | 11 | 1119 | 19 | 10487958 | 1169 |  | 2150
| 20 |  |  | 0 | 70 | 70 | 514531 | 1 | 120 | 20 | 990161 | 1 | 170 | 70 | 1465634 | 2 | 220 | 20 | 1940282 | 700 | 2 | 270 | 70 | 2414930 | 3 | 320 | 20 | 2890230 | 3 | 370 | 70 | 3366003 | 4 | 420 | 20 | 3841776 | 4 | 470 | 70 | 4317045 | 700 | 5 | 520 | 20 | 4791812 | 5 | 570 | 70 | 5266612 | 6 | 620 | 20 | 5741791 | 6 | 670 | 70 | 6216969 | 7 | 720 | 20 | 6692147 | 1340 | 7 | 770 | 70 | 7167415 | 8 | 820 | 20 | 7642828 | 8 | 870 | 70 | 8118242 | 9 | 920 | 20 | 8593655 | 9 | 970 | 70 | 9069487 | 10 | 1020 | 20 | 9545482 | 10 | 1070 | 70 | 10021476 | 11 | 1120 | 20 | 10497484 | 1170 |  | 2150
| 21 |  |  | 0 | 71 | 71 | 524045 | 1 | 121 | 21 | 999671 | 1 | 171 | 71 | 1475127 | 2 | 221 | 21 | 1949775 | 700 | 2 | 271 | 71 | 2424423 | 3 | 321 | 21 | 2899746 | 3 | 371 | 71 | 3375518 | 4 | 421 | 21 | 3851291 | 4 | 471 | 71 | 4326541 | 700 | 5 | 521 | 21 | 4801308 | 5 | 571 | 71 | 5276116 | 6 | 621 | 21 | 5751294 | 6 | 671 | 71 | 6226473 | 7 | 721 | 21 | 6701651 | 1340 | 7 | 771 | 71 | 7176923 | 8 | 821 | 21 | 7652337 | 8 | 871 | 71 | 8127750 | 9 | 921 | 21 | 8603164 | 9 | 971 | 71 | 9079007 | 10 | 1021 | 21 | 9555002 | 10 | 1071 | 71 | 10030996 | 11 | 1121 | 21 | 10507011 | 1171 |  | 2150
| 22 |  |  | 0 | 72 | 72 | 533559 | 1 | 122 | 22 | 1009181 | 1 | 172 | 72 | 1484620 | 2 | 222 | 22 | 1959268 | 700 | 2 | 272 | 72 | 2433916 | 3 | 322 | 22 | 2909261 | 3 | 372 | 72 | 3385034 | 4 | 422 | 22 | 3860807 | 4 | 472 | 72 | 4336036 | 700 | 5 | 522 | 22 | 4810803 | 5 | 572 | 72 | 5285619 | 6 | 622 | 22 | 5760798 | 6 | 672 | 72 | 6235976 | 7 | 722 | 22 | 6711155 | 1340 | 7 | 772 | 72 | 7186432 | 8 | 822 | 22 | 7661845 | 8 | 872 | 72 | 8137258 | 9 | 922 | 22 | 8612672 | 9 | 972 | 72 | 9088527 | 10 | 1022 | 22 | 9564521 | 10 | 1072 | 72 | 10040516 | 11 | 1122 | 22 | 10516537 | 1172 |  | 2150
| 23 |  |  | 0 | 73 | 73 | 543072 | 1 | 123 | 23 | 1018691 | 1 | 173 | 73 | 1494113 | 2 | 223 | 23 | 1968761 | 700 | 2 | 273 | 73 | 2443409 | 3 | 323 | 23 | 2918777 | 3 | 373 | 73 | 3394549 | 4 | 423 | 23 | 3870322 | 4 | 473 | 73 | 4345531 | 700 | 5 | 523 | 23 | 4820298 | 5 | 573 | 73 | 5295123 | 6 | 623 | 23 | 5770301 | 6 | 673 | 73 | 6245480 | 7 | 723 | 23 | 6720658 | 1340 | 7 | 773 | 73 | 7195940 | 8 | 823 | 23 | 7671353 | 8 | 873 | 73 | 8146767 | 9 | 923 | 23 | 8622180 | 9 | 973 | 73 | 9098047 | 10 | 1023 | 23 | 9574041 | 10 | 1073 | 73 | 10050036 | 11 | 1123 | 23 | 10526064 | 1173 |  | 2150
| 24 |  |  | 0 | 74 | 74 | 552586 | 1 | 124 | 24 | 1028202 | 1 | 174 | 74 | 1503606 | 2 | 224 | 24 | 1978254 | 700 | 2 | 274 | 74 | 2452902 | 3 | 324 | 24 | 2928292 | 3 | 374 | 74 | 3404065 | 4 | 424 | 24 | 3879837 | 4 | 474 | 74 | 4355027 | 700 | 5 | 524 | 24 | 4829794 | 5 | 574 | 74 | 5304627 | 6 | 624 | 24 | 5779805 | 6 | 674 | 74 | 6254983 | 7 | 724 | 24 | 6730162 | 1340 | 7 | 774 | 74 | 7205448 | 8 | 824 | 24 | 7680862 | 8 | 874 | 74 | 8156275 | 9 | 924 | 24 | 8631688 | 9 | 974 | 74 | 9107567 | 10 | 1024 | 24 | 9583561 | 10 | 1074 | 74 | 10059556 | 11 | 1124 | 24 | 10535590 | 1174 |  | 2150
| 25 |  |  | 0 | 75 | 75 | 562100 | 1 | 125 | 25 | 1037712 | 1 | 175 | 75 | 1513099 | 2 | 225 | 25 | 1987747 | 700 | 2 | 275 | 75 | 2462395 | 3 | 325 | 25 | 2937807 | 3 | 375 | 75 | 3413580 | 4 | 425 | 25 | 3889353 | 4 | 475 | 75 | 4364522 | 700 | 5 | 525 | 25 | 4839289 | 5 | 575 | 75 | 5314130 | 6 | 625 | 25 | 5789308 | 6 | 675 | 75 | 6264487 | 7 | 725 | 25 | 6739665 | 1340 | 7 | 775 | 75 | 7214956 | 8 | 825 | 25 | 7690370 | 8 | 875 | 75 | 8165783 | 9 | 925 | 25 | 8641197 | 9 | 975 | 75 | 9117087 | 10 | 1025 | 25 | 9593081 | 10 | 1075 | 75 | 10069076 | 11 | 1125 | 25 | 10545117 | 1175 |  | 2150
| 26 |  |  | 0 | 76 | 76 | 571613 | 1 | 126 | 26 | 1047222 | 1 | 176 | 76 | 1522592 | 2 | 226 | 26 | 1997240 | 700 | 2 | 276 | 76 | 2471888 | 3 | 326 | 26 | 2947323 | 3 | 376 | 76 | 3423096 | 4 | 426 | 26 | 3898868 | 4 | 476 | 76 | 4374018 | 700 | 5 | 526 | 26 | 4848784 | 5 | 576 | 76 | 5323634 | 6 | 626 | 26 | 5798812 | 6 | 676 | 76 | 6273990 | 7 | 726 | 26 | 6749169 | 1340 | 7 | 776 | 76 | 7224465 | 8 | 826 | 26 | 7699878 | 8 | 876 | 76 | 8175291 | 9 | 926 | 26 | 8650705 | 9 | 976 | 76 | 9126606 | 10 | 1026 | 26 | 9602601 | 10 | 1076 | 76 | 10078596 | 11 | 1126 | 26 | 10554643 | 1176 |  | 2150
| 27 |  |  | 0 | 77 | 77 | 581127 | 1 | 127 | 27 | 1056732 | 1 | 177 | 77 | 1532085 | 2 | 227 | 27 | 2006733 | 700 | 2 | 277 | 77 | 2481381 | 3 | 327 | 27 | 2956838 | 3 | 377 | 77 | 3432611 | 4 | 427 | 27 | 3908384 | 4 | 477 | 77 | 4383513 | 700 | 5 | 527 | 27 | 4858280 | 5 | 577 | 77 | 5333137 | 6 | 627 | 27 | 5808316 | 6 | 677 | 77 | 6283494 | 7 | 727 | 27 | 6758672 | 1340 | 7 | 777 | 77 | 7233973 | 8 | 827 | 27 | 7709386 | 8 | 877 | 77 | 8184800 | 9 | 927 | 27 | 8660213 | 9 | 977 | 77 | 9136126 | 10 | 1027 | 27 | 9612121 | 10 | 1077 | 77 | 10088115 | 11 | 1127 | 27 | 10564170 | 1177 |  | 2150
| 28 |  |  | 0 | 78 | 78 | 590640 | 1 | 128 | 28 | 1066242 | 1 | 178 | 78 | 1541578 | 2 | 228 | 28 | 2016226 | 700 | 2 | 278 | 78 | 2490874 | 3 | 328 | 28 | 2966354 | 3 | 378 | 78 | 3442127 | 4 | 428 | 28 | 3917899 | 4 | 478 | 78 | 4393008 | 700 | 5 | 528 | 28 | 4867775 | 5 | 578 | 78 | 5342641 | 6 | 628 | 28 | 5817819 | 6 | 678 | 78 | 6292998 | 7 | 728 | 28 | 6768176 | 1340 | 7 | 778 | 78 | 7243481 | 8 | 828 | 28 | 7718895 | 8 | 878 | 78 | 8194308 | 9 | 928 | 28 | 8669721 | 9 | 978 | 78 | 9145646 | 10 | 1028 | 28 | 9621641 | 10 | 1078 | 78 | 10097635 | 11 | 1128 | 28 | 10573696 | 1178 |  | 2150
| 29 |  |  | 0 | 79 | 79 | 600154 | 1 | 129 | 29 | 1075752 | 1 | 179 | 79 | 1551071 | 2 | 229 | 29 | 2025719 | 700 | 2 | 279 | 79 | 2500366 | 3 | 329 | 29 | 2975869 | 3 | 379 | 79 | 3451642 | 4 | 429 | 29 | 3927415 | 4 | 479 | 79 | 4402504 | 700 | 5 | 529 | 29 | 4877270 | 5 | 579 | 79 | 5352144 | 6 | 629 | 29 | 5827323 | 6 | 679 | 79 | 6302501 | 7 | 729 | 29 | 6777680 | 1340 | 7 | 779 | 79 | 7252989 | 8 | 829 | 29 | 7728403 | 8 | 879 | 79 | 8203816 | 9 | 929 | 29 | 8679230 | 9 | 979 | 79 | 9155166 | 10 | 1029 | 29 | 9631161 | 10 | 1079 | 79 | 10107155 | 11 | 1129 | 29 | 10583223 | 1179 |  | 2150
| 30 |  |  | 0 | 80 | 80 | 609668 | 1 | 130 | 30 | 1085262 | 1 | 180 | 80 | 1560564 | 2 | 230 | 30 | 2035211 | 700 | 2 | 280 | 80 | 2509859 | 3 | 330 | 30 | 2985385 | 3 | 380 | 80 | 3461157 | 4 | 430 | 30 | 3936930 | 4 | 480 | 80 | 4411999 | 700 | 5 | 530 | 30 | 4886766 | 5 | 580 | 80 | 5361648 | 6 | 630 | 30 | 5836826 | 6 | 680 | 80 | 6312005 | 7 | 730 | 30 | 6787183 | 1340 | 7 | 780 | 80 | 7262498 | 8 | 830 | 30 | 7737911 | 8 | 880 | 80 | 8213325 | 9 | 930 | 30 | 8688738 | 9 | 980 | 80 | 9164686 | 10 | 1030 | 30 | 9640681 | 10 | 1080 | 80 | 10116675 | 11 | 1130 | 30 | 10592749 | 1180 |  | 2150
| 31 |  |  | 0 | 81 | 81 | 619182 | 1 | 131 | 31 | 1094773 | 1 | 181 | 81 | 1570056 | 2 | 231 | 31 | 2044704 | 700 | 2 | 281 | 81 | 2519352 | 3 | 331 | 31 | 2994900 | 3 | 381 | 81 | 3470673 | 4 | 431 | 31 | 3946446 | 4 | 481 | 81 | 4421494 | 700 | 5 | 531 | 31 | 4896261 | 5 | 581 | 81 | 5371152 | 6 | 631 | 31 | 5846330 | 6 | 681 | 81 | 6321508 | 7 | 731 | 31 | 6796687 | 1340 | 7 | 781 | 81 | 7272006 | 8 | 831 | 31 | 7747419 | 8 | 881 | 81 | 8222833 | 9 | 931 | 31 | 8698246 | 9 | 981 | 81 | 9174206 | 10 | 1031 | 31 | 9650200 | 10 | 1081 | 81 | 10126195 | 11 | 1131 | 31 | 10602276 | 1181 |  | 2150
| 32 |  |  | 0 | 82 | 82 | 628696 | 1 | 132 | 32 | 1104283 | 1 | 182 | 82 | 1579549 | 2 | 232 | 32 | 2054197 | 700 | 2 | 282 | 82 | 2528845 | 3 | 332 | 32 | 3004416 | 3 | 382 | 82 | 3480188 | 4 | 432 | 32 | 3955961 | 4 | 482 | 82 | 4430990 | 700 | 5 | 532 | 32 | 4905756 | 5 | 582 | 82 | 5380655 | 6 | 632 | 32 | 5855833 | 6 | 682 | 82 | 6331012 | 7 | 732 | 32 | 6806190 | 1340 | 7 | 782 | 82 | 7281514 | 8 | 832 | 32 | 7756928 | 8 | 882 | 82 | 8232341 | 9 | 932 | 32 | 8707754 | 9 | 982 | 82 | 9183726 | 10 | 1032 | 32 | 9659720 | 10 | 1082 | 82 | 10135715 | 11 | 1132 | 32 | 10611802 | 1182 |  | 2150
| 33 |  |  | 0 | 83 | 83 | 638209 | 1 | 133 | 33 | 1113793 | 1 | 183 | 83 | 1589042 | 2 | 233 | 33 | 2063690 | 700 | 2 | 283 | 83 | 2538338 | 3 | 333 | 33 | 3013931 | 3 | 383 | 83 | 3489704 | 4 | 433 | 33 | 3965477 | 4 | 483 | 83 | 4440485 | 700 | 5 | 533 | 33 | 4915252 | 5 | 583 | 83 | 5390159 | 6 | 633 | 33 | 5865337 | 6 | 683 | 83 | 6340515 | 7 | 733 | 33 | 6815694 | 1340 | 7 | 783 | 83 | 7291023 | 8 | 833 | 33 | 7766436 | 8 | 883 | 83 | 8241849 | 9 | 933 | 33 | 8717263 | 9 | 983 | 83 | 9193246 | 10 | 1033 | 33 | 9669240 | 10 | 1083 | 83 | 10145235 | 11 | 1133 | 33 | 10621329 | 1183 |  | 2150
| 34 |  |  | 0 | 84 | 84 | 647723 | 1 | 134 | 34 | 1123303 | 1 | 184 | 84 | 1598535 | 2 | 234 | 34 | 2073183 | 700 | 2 | 284 | 84 | 2547831 | 3 | 334 | 34 | 3023447 | 3 | 384 | 84 | 3499219 | 4 | 434 | 34 | 3974992 | 4 | 484 | 84 | 4449980 | 700 | 5 | 534 | 34 | 4924747 | 5 | 584 | 84 | 5399662 | 6 | 634 | 34 | 5874841 | 6 | 684 | 84 | 6350019 | 7 | 734 | 34 | 6825197 | 1340 | 7 | 784 | 84 | 7300531 | 8 | 834 | 34 | 7775944 | 8 | 884 | 84 | 8251358 | 9 | 934 | 34 | 8726771 | 9 | 984 | 84 | 9202766 | 10 | 1034 | 34 | 9678760 | 10 | 1084 | 84 | 10154755 | 11 | 1134 | 34 | 10630855 | 1184 |  | 2150
| 35 |  |  | 0 | 85 | 85 | 657237 | 1 | 135 | 35 | 1132813 | 1 | 185 | 85 | 1608028 | 2 | 235 | 35 | 2082676 | 700 | 2 | 285 | 85 | 2557324 | 3 | 335 | 35 | 3032962 | 3 | 385 | 85 | 3508735 | 4 | 435 | 35 | 3984507 | 4 | 485 | 85 | 4459476 | 700 | 5 | 535 | 35 | 4934242 | 5 | 585 | 85 | 5409166 | 6 | 635 | 35 | 5884344 | 6 | 685 | 85 | 6359523 | 7 | 735 | 35 | 6834701 | 1340 | 7 | 785 | 85 | 7310039 | 8 | 835 | 35 | 7785452 | 8 | 885 | 85 | 8260866 | 9 | 935 | 35 | 8736291 | 9 | 985 | 85 | 9212285 | 10 | 1035 | 35 | 9688280 | 10 | 1085 | 85 | 10164275 | 11 | 1135 | 35 | 10640382 | 1185 |  | 2150
| 36 |  |  | 0 | 86 | 86 | 666751 | 1 | 136 | 36 | 1142323 | 1 | 186 | 86 | 1617521 | 2 | 236 | 36 | 2092169 | 700 | 2 | 286 | 86 | 2566817 | 3 | 336 | 36 | 3042477 | 3 | 386 | 86 | 3518250 | 4 | 436 | 36 | 3994023 | 4 | 486 | 86 | 4468971 | 700 | 5 | 536 | 36 | 4943738 | 5 | 586 | 86 | 5418669 | 6 | 636 | 36 | 5893848 | 6 | 686 | 86 | 6369026 | 7 | 736 | 36 | 6844204 | 1340 | 7 | 786 | 86 | 7319547 | 8 | 836 | 36 | 7794961 | 8 | 886 | 86 | 8270374 | 9 | 936 | 36 | 8745811 | 9 | 986 | 86 | 9221805 | 10 | 1036 | 36 | 9697800 | 10 | 1086 | 86 | 10173794 | 11 | 1136 | 36 | 10649908 | 1186 |  | 2150
| 37 |  |  | 0 | 87 | 87 | 676266 | 1 | 137 | 37 | 1151834 | 1 | 187 | 87 | 1627014 | 2 | 237 | 37 | 2101662 | 700 | 2 | 287 | 87 | 2576310 | 3 | 337 | 37 | 3051993 | 3 | 387 | 87 | 3527766 | 4 | 437 | 37 | 4003538 | 4 | 487 | 87 | 4478466 | 700 | 5 | 537 | 37 | 4953233 | 5 | 587 | 87 | 5428173 | 6 | 637 | 37 | 5903351 | 6 | 687 | 87 | 6378530 | 7 | 737 | 37 | 6853708 | 1340 | 7 | 787 | 87 | 7329056 | 8 | 837 | 37 | 7804469 | 8 | 887 | 87 | 8279882 | 9 | 937 | 37 | 8755331 | 9 | 987 | 87 | 9231325 | 10 | 1037 | 37 | 9707320 | 10 | 1087 | 87 | 10183314 | 11 | 1137 | 37 | 10659435 | 1187 |  | 2150
| 38 |  |  | 0 | 88 | 88 | 685781 | 1 | 138 | 38 | 1161344 | 1 | 188 | 88 | 1636507 | 2 | 238 | 38 | 2111155 | 700 | 2 | 288 | 88 | 2585803 | 3 | 338 | 38 | 3061508 | 3 | 388 | 88 | 3537281 | 4 | 438 | 38 | 4013054 | 4 | 488 | 88 | 4487962 | 700 | 5 | 538 | 38 | 4962729 | 5 | 588 | 88 | 5437676 | 6 | 638 | 38 | 5912855 | 6 | 688 | 88 | 6388033 | 7 | 738 | 38 | 6863212 | 1340 | 7 | 788 | 88 | 7338564 | 8 | 838 | 38 | 7813977 | 8 | 888 | 88 | 8289391 | 9 | 938 | 38 | 8764851 | 9 | 988 | 88 | 9240845 | 10 | 1038 | 38 | 9716840 | 10 | 1088 | 88 | 10192834 | 11 | 1138 | 38 | 10668961 | 1188 |  | 2150
| 39 |  |  | 0 | 89 | 89 | 695296 | 1 | 139 | 39 | 1170854 | 1 | 189 | 89 | 1646000 | 2 | 239 | 39 | 2120648 | 700 | 2 | 289 | 89 | 2595296 | 3 | 339 | 39 | 3071024 | 3 | 389 | 89 | 3546797 | 4 | 439 | 39 | 4022569 | 4 | 489 | 89 | 4497457 | 700 | 5 | 539 | 39 | 4972224 | 5 | 589 | 89 | 5447180 | 6 | 639 | 39 | 5922358 | 6 | 689 | 89 | 6397537 | 7 | 739 | 39 | 6872715 | 1340 | 7 | 789 | 89 | 7348072 | 8 | 839 | 39 | 7823486 | 8 | 889 | 89 | 8298899 | 9 | 939 | 39 | 8774370 | 9 | 989 | 89 | 9250365 | 10 | 1039 | 39 | 9726360 | 10 | 1089 | 89 | 10202354 | 11 | 1139 | 39 | 10678488 | 1189 |  | 2150
| 40 |  |  | 0 | 90 | 90 | 704812 | 1 | 140 | 40 | 1180364 | 1 | 190 | 90 | 1655493 | 2 | 240 | 40 | 2130141 | 700 | 2 | 290 | 90 | 2604789 | 3 | 340 | 40 | 3080539 | 3 | 390 | 90 | 3556312 | 4 | 440 | 40 | 4032085 | 4 | 490 | 90 | 4506952 | 700 | 5 | 540 | 40 | 4981719 | 5 | 590 | 90 | 5456684 | 6 | 640 | 40 | 5931862 | 6 | 690 | 90 | 6407040 | 7 | 740 | 40 | 6882219 | 1340 | 7 | 790 | 90 | 7357580 | 8 | 840 | 40 | 7832994 | 8 | 890 | 90 | 8308407 | 9 | 940 | 40 | 8783890 | 9 | 990 | 90 | 9259885 | 10 | 1040 | 40 | 9735879 | 10 | 1090 | 90 | 10211874 | 11 | 1140 | 40 | 10688014 | 1190 |  | 2150
| 41 |  |  | 0 | 91 | 91 | 714327 | 1 | 141 | 41 | 1189874 | 1 | 191 | 91 | 1664986 | 2 | 241 | 41 | 2139634 | 700 | 2 | 291 | 91 | 2614282 | 3 | 341 | 41 | 3090055 | 3 | 391 | 91 | 3565827 | 4 | 441 | 41 | 4041600 | 4 | 491 | 91 | 4516448 | 700 | 5 | 541 | 41 | 4991215 | 5 | 591 | 91 | 5466187 | 6 | 641 | 41 | 5941366 | 6 | 691 | 91 | 6416544 | 7 | 741 | 41 | 6891722 | 1340 | 7 | 791 | 91 | 7367089 | 8 | 841 | 41 | 7842502 | 8 | 891 | 91 | 8317915 | 9 | 941 | 41 | 8793410 | 9 | 991 | 91 | 9269405 | 10 | 1041 | 41 | 9745399 | 10 | 1091 | 91 | 10221394 | 11 | 1141 | 41 | 10697541 | 1191 |  | 2150
| 42 |  |  | 0 | 92 | 92 | 723842 | 1 | 142 | 42 | 1199384 | 1 | 192 | 92 | 1674479 | 2 | 242 | 42 | 2149127 | 700 | 2 | 292 | 92 | 2623797 | 3 | 342 | 42 | 3099570 | 3 | 392 | 92 | 3575343 | 4 | 442 | 42 | 4051116 | 4 | 492 | 92 | 4525943 | 700 | 5 | 542 | 42 | 5000710 | 5 | 592 | 92 | 5475691 | 6 | 642 | 42 | 5950869 | 6 | 692 | 92 | 6426048 | 7 | 742 | 42 | 6901226 | 1340 | 7 | 792 | 92 | 7376597 | 8 | 842 | 42 | 7852010 | 8 | 892 | 92 | 8327424 | 9 | 942 | 42 | 8802930 | 9 | 992 | 92 | 9278925 | 10 | 1042 | 42 | 9754919 | 10 | 1092 | 92 | 10230914 | 11 | 1142 | 42 | 10707067 | 1192 |  | 2150
| 43 |  |  | 0 | 93 | 93 | 733357 | 1 | 143 | 43 | 1208894 | 1 | 193 | 93 | 1683972 | 2 | 243 | 43 | 2158620 | 700 | 2 | 293 | 93 | 2633313 | 3 | 343 | 43 | 3109086 | 3 | 393 | 93 | 3584858 | 4 | 443 | 43 | 4060631 | 4 | 493 | 93 | 4535438 | 700 | 5 | 543 | 43 | 5010205 | 5 | 593 | 93 | 5485194 | 6 | 643 | 43 | 5960373 | 6 | 693 | 93 | 6435551 | 7 | 743 | 43 | 6910729 | 1340 | 7 | 793 | 93 | 7386105 | 8 | 843 | 43 | 7861519 | 8 | 893 | 93 | 8336932 | 9 | 943 | 43 | 8812450 | 9 | 993 | 93 | 9288445 | 10 | 1043 | 43 | 9764439 | 10 | 1093 | 93 | 10240434 | 11 | 1143 | 43 | 10716594 | 1193 |  | 2150
| 44 |  |  | 0 | 94 | 94 | 742872 | 1 | 144 | 44 | 1218405 | 1 | 194 | 94 | 1693465 | 2 | 244 | 44 | 2168113 | 700 | 2 | 294 | 94 | 2642828 | 3 | 344 | 44 | 3118601 | 3 | 394 | 94 | 3594374 | 4 | 444 | 44 | 4070147 | 4 | 494 | 94 | 4544934 | 700 | 5 | 544 | 44 | 5019701 | 5 | 594 | 94 | 5494698 | 6 | 644 | 44 | 5969876 | 6 | 694 | 94 | 6445055 | 7 | 744 | 44 | 6920233 | 1340 | 7 | 794 | 94 | 7395614 | 8 | 844 | 44 | 7871027 | 8 | 894 | 94 | 8346440 | 9 | 944 | 44 | 8821970 | 9 | 994 | 94 | 9297964 | 10 | 1044 | 44 | 9773959 | 10 | 1094 | 94 | 10249954 | 11 | 1144 | 44 | 10726120 | 1194 |  | 2150
| 45 |  |  | 0 | 95 | 95 | 752387 | 1 | 145 | 45 | 1227915 | 1 | 195 | 95 | 1702958 | 2 | 245 | 45 | 2177606 | 700 | 2 | 295 | 95 | 2652344 | 3 | 345 | 45 | 3128117 | 3 | 395 | 95 | 3603889 | 4 | 445 | 45 | 4079662 | 4 | 495 | 95 | 4554429 | 700 | 5 | 545 | 45 | 5029196 | 5 | 595 | 95 | 5504201 | 6 | 645 | 45 | 5979380 | 6 | 695 | 95 | 6454558 | 7 | 745 | 45 | 6929737 | 1340 | 7 | 795 | 95 | 7405122 | 8 | 845 | 45 | 7880535 | 8 | 895 | 95 | 8355949 | 9 | 945 | 45 | 8831490 | 9 | 995 | 95 | 9307484 | 10 | 1045 | 45 | 9783479 | 10 | 1095 | 95 | 10259474 | 11 | 1145 | 45 | 10735647 | 1195 |  | 2150
| 46 |  |  | 0 | 96 | 96 | 761903 | 1 | 146 | 46 | 1237425 | 1 | 196 | 96 | 1712451 | 2 | 246 | 46 | 2187099 | 700 | 2 | 296 | 96 | 2661859 | 3 | 346 | 46 | 3137632 | 3 | 396 | 96 | 3613405 | 4 | 446 | 46 | 4089157 | 4 | 496 | 96 | 4563924 | 700 | 5 | 546 | 46 | 5038691 | 5 | 596 | 96 | 5513705 | 6 | 646 | 46 | 5988883 | 6 | 696 | 96 | 6464062 | 7 | 746 | 46 | 6939240 | 1340 | 7 | 796 | 96 | 7414630 | 8 | 846 | 46 | 7890043 | 8 | 896 | 96 | 8365457 | 9 | 946 | 46 | 8841010 | 9 | 996 | 96 | 9317004 | 10 | 1046 | 46 | 9792999 | 10 | 1096 | 96 | 10268993 | 11 | 1146 | 46 | 10745173 | 1196 |  | 2150
| 47 |  |  | 0 | 97 | 97 | 771418 | 1 | 147 | 47 | 1246935 | 1 | 197 | 97 | 1721944 | 2 | 247 | 47 | 2196592 | 700 | 2 | 297 | 97 | 2671375 | 3 | 347 | 47 | 3147147 | 3 | 397 | 97 | 3622920 | 4 | 447 | 47 | 4098653 | 4 | 497 | 97 | 4573420 | 700 | 5 | 547 | 47 | 5048187 | 5 | 597 | 97 | 5523209 | 6 | 647 | 47 | 5998387 | 6 | 697 | 97 | 6473565 | 7 | 747 | 47 | 6948744 | 1340 | 7 | 797 | 97 | 7424138 | 8 | 847 | 47 | 7899552 | 8 | 897 | 97 | 8374965 | 9 | 947 | 47 | 8850530 | 9 | 997 | 97 | 9326524 | 10 | 1047 | 47 | 9802519 | 10 | 1097 | 97 | 10278513 | 11 | 1147 | 47 | 10754700 | 1197 |  | 2150
| 48 |  |  | 0 | 98 | 98 | 780933 | 1 | 148 | 48 | 1256445 | 1 | 198 | 98 | 1731437 | 2 | 248 | 48 | 2206085 | 700 | 2 | 298 | 98 | 2680890 | 3 | 348 | 48 | 3156663 | 3 | 398 | 98 | 3632436 | 4 | 448 | 48 | 4108148 | 4 | 498 | 98 | 4582915 | 700 | 5 | 548 | 48 | 5057682 | 5 | 598 | 98 | 5532712 | 6 | 648 | 48 | 6007891 | 6 | 698 | 98 | 6483069 | 7 | 748 | 48 | 6958247 | 1340 | 7 | 798 | 98 | 7433647 | 8 | 848 | 48 | 7909060 | 8 | 898 | 98 | 8384473 | 9 | 948 | 48 | 8860049 | 9 | 998 | 98 | 9336044 | 10 | 1048 | 48 | 9812039 | 10 | 1098 | 98 | 10288033 | 11 | 1148 | 48 | 10764226 | 1198 |  | 2150
| 49 |  |  | 0 | 99 | 99 | 790448 | 1 | 149 | 49 | 1265955 | 1 | 199 | 99 | 1740930 | 2 | 249 | 49 | 2215578 | 700 | 2 | 299 | 99 | 2690406 | 3 | 349 | 49 | 3166178 | 3 | 399 | 99 | 3641951 | 4 | 449 | 49 | 4117643 | 4 | 499 | 99 | 4592410 | 700 | 5 | 549 | 49 | 5067177 | 5 | 599 | 99 | 5542216 | 6 | 649 | 49 | 6017394 | 6 | 699 | 99 | 6492572 | 7 | 749 | 49 | 6967751 | 1340 | 7 | 799 | 99 | 7443155 | 8 | 849 | 49 | 7918568 | 8 | 899 | 99 | 8393982 | 9 | 949 | 49 | 8869569 | 9 | 999 | 99 | 9345564 | 10 | 1049 | 49 | 9821559 | 10 | 1099 | 99 | 10297553 | 11 | 1149 | 49 | 10773753 | 1199 |  | 2150
0 | 50 | 50 | 324259 | 1 | 100 | 0 | 799958 | 1 | 150 | 50 | 1275465 | 2 | 200 | 0 | 1750423 | 2 | 250 | 50 | 2225071 | 700 | 3 | 300 | 0 | 2699921 | 3 | 350 | 50 | 3175694 | 4 | 400 | 0 | 3651467 | 4 | 450 | 50 | 4127139 | 5 | 500 | 0 | 4601906 | 700 | 5 | 550 | 50 | 5076673 | 6 | 600 | 0 | 5551719 | 6 | 650 | 50 | 6026898 | 7 | 700 | 0 | 6502076 | 7 | 750 | 50 | 6977254 | 1340 | 8 | 800 | 0 | 7452663 | 8 | 850 | 50 | 7928077 | 9 | 900 | 0 | 8403490 | 9 | 950 | 50 | 8879089 | 10 | 1000 | 0 | 9355084 | 10 | 1050 | 50 | 9831078 | 11 | 1100 | 0 | 10307073 | 11 | 1150 | 50 | 10783279 | 1200 |  | 2150

## Header Tank (hidden)
Lαmpirαn-4: TKO
FORMULIR PEMELIHARAAN
TANK MASTER DATA & STRAPPING TABLE
TANGGAL | :
NO. SURAT | :
Creαte Tαnk Mαster Dαtα / Strαpping Tαble
Chαnge Tαnk Mαster Dαtα / Strαpping Tαble
PLANT | : | R501
JUMLAH LAMPIRAN | : | 3 (tigα) lembαr
DATA TANKI
Tαnk Number | : | T00902 | Code Product | : | AU
Type | : | FIXED ROOF | Terα Exprired | : | Oktober 2029
Areα | : | SOUTH | Lαst Terα Dαte | : | 2019-10-11
Unit | : | Balikpapan - Kalimantan Timur | Cαlibrαtion Stαtus | :
Stαtus Tαnk | : | Bαik / Aktif Digunαkαn | Lαst Cαlibrαte By | : | 1. Hikmat Rijadi
Product | : | ARJUNA CRUDE |  |  | 2. Anditya Kristianto
Diαmeter | : | 34820 |  |  |  | TAG PI:
Tαnk Height | : | 11818 | Min. Tinggi Cαirαn | : | 500 | Tαg PI Pressure | :
Totαl Tαnk Height | : | 11818 | Mαx. Tinggi Cαirαn | : | 11500 | Tαg PI Temperαture | :
Dαtum Dip | : | 82 | Fαktor Koreksi Vol | : | 36 | Tαg PI Level | :
Dip Hole | : | 12671 | Vαlve Fαctor | : | TIDAK ADA
Drαwαble Height | : | 500 | Pαint Fαctor | : | TIDAK ADA
Mαssα Jenis | : | 0.86 | Clinggαge Fαctor | : | TIDAK ADA
Suhu | : | 36 | Koefisien Muαi | : | 3.48e-05
Volume Bersih | : | 10783279 | Koefiesien Jenis Minyαk | :
|  |  | Koefisien Type Atαp | : | TIDAK ADA
|  | Pejabat yang menyetujui |  |  |  | User yang meminta
|  | Arief Maukar R. |  |  |  | Widy Nungroho
|  | 733433 |  |  |  | 745124

## Strapping Tank (hidden)
Lαmpirαn-4: TKO
STRAPPING TABLE
Plαnt |  | : | R501
Tαnk Number |  | : | T00902
Lαmpirαn Ke |  | :
Perhitungαn Mejα Ukur |  | : | TIDAK ADA
Height (cm) | Volume (liter)
50 | 324259
51 | 333773
52 | 343286
53 | 352800
54 | 362313
55 | 371827
56 | 381341
57 | 390854
58 | 400368
59 | 409882
60 | 419395
61 | 428909
62 | 438422
63 | 447936
64 | 457450
65 | 466963
66 | 476477
67 | 485991
68 | 495504
69 | 505018
70 | 514531
71 | 524045
72 | 533559
73 | 543072
74 | 552586
75 | 562100
76 | 571613
77 | 581127
78 | 590640
79 | 600154
80 | 609668
81 | 619182
82 | 628696
83 | 638209
84 | 647723
85 | 657237
86 | 666751
87 | 676266
88 | 685781
89 | 695296
90 | 704812
91 | 714327
92 | 723842
93 | 733357
94 | 742872
95 | 752387
96 | 761903
97 | 771418
98 | 780933
99 | 790448
100 | 799958
101 | 809468
102 | 818978
103 | 828489
104 | 837999
105 | 847509
106 | 857019
107 | 866529
108 | 876039
109 | 885549
110 | 895060
111 | 904570
112 | 914080
113 | 923590
114 | 933100
115 | 942610
116 | 952120
117 | 961631
118 | 971141
119 | 980651
120 | 990161
121 | 999671
122 | 1009181
123 | 1018691
124 | 1028202
125 | 1037712
126 | 1047222
127 | 1056732
128 | 1066242
129 | 1075752
130 | 1085262
131 | 1094773
132 | 1104283
133 | 1113793
134 | 1123303
135 | 1132813
136 | 1142323
137 | 1151834
138 | 1161344
139 | 1170854
140 | 1180364
141 | 1189874
142 | 1199384
143 | 1208894
144 | 1218405
145 | 1227915
146 | 1237425
147 | 1246935
148 | 1256445
149 | 1265955
150 | 1275465
151 | 1284976
152 | 1294486
153 | 1303996
154 | 1313506
155 | 1323016
156 | 1332526
157 | 1342036
158 | 1351547
159 | 1361057
160 | 1370567
161 | 1380077
162 | 1389587
163 | 1399097
164 | 1408607
165 | 1418118
166 | 1427628
167 | 1437138
168 | 1446648
169 | 1456141
170 | 1465634
171 | 1475127
172 | 1484620
173 | 1494113
174 | 1503606
175 | 1513099
176 | 1522592
177 | 1532085
178 | 1541578
179 | 1551071
180 | 1560564
181 | 1570056
182 | 1579549
183 | 1589042
184 | 1598535
185 | 1608028
186 | 1617521
187 | 1627014
188 | 1636507
189 | 1646000
190 | 1655493
191 | 1664986
192 | 1674479
193 | 1683972
194 | 1693465
195 | 1702958
196 | 1712451
197 | 1721944
198 | 1731437
199 | 1740930
200 | 1750423
201 | 1759916
202 | 1769409
203 | 1778902
204 | 1788395
205 | 1797887
206 | 1807380
207 | 1816873
208 | 1826366
209 | 1835859
210 | 1845352
211 | 1854845
212 | 1864338
213 | 1873831
214 | 1883324
215 | 1892817
216 | 1902310
217 | 1911803
218 | 1921296
219 | 1930789
220 | 1940282
221 | 1949775
222 | 1959268
223 | 1968761
224 | 1978254
225 | 1987747
226 | 1997240
227 | 2006733
228 | 2016226
229 | 2025719
230 | 2035211
231 | 2044704
232 | 2054197
233 | 2063690
234 | 2073183
235 | 2082676
236 | 2092169
237 | 2101662
238 | 2111155
239 | 2120648
240 | 2130141
241 | 2139634
242 | 2149127
243 | 2158620
244 | 2168113
245 | 2177606
246 | 2187099
247 | 2196592
248 | 2206085
249 | 2215578
250 | 2225071
251 | 2234564
252 | 2244057
253 | 2253550
254 | 2263043
255 | 2272535
256 | 2282028
257 | 2291521
258 | 2301014
259 | 2310507
260 | 2320000
261 | 2329493
262 | 2338986
263 | 2348479
264 | 2357972
265 | 2367465
266 | 2376958
267 | 2386451
268 | 2395944
269 | 2405437
270 | 2414930
271 | 2424423
272 | 2433916
273 | 2443409
274 | 2452902
275 | 2462395
276 | 2471888
277 | 2481381
278 | 2490874
279 | 2500366
280 | 2509859
281 | 2519352
282 | 2528845
283 | 2538338
284 | 2547831
285 | 2557324
286 | 2566817
287 | 2576310
288 | 2585803
289 | 2595296
290 | 2604789
291 | 2614282
292 | 2623797
293 | 2633313
294 | 2642828
295 | 2652344
296 | 2661859
297 | 2671375
298 | 2680890
299 | 2690406
300 | 2699921
301 | 2709437
302 | 2718952
303 | 2728467
304 | 2737983
305 | 2747498
306 | 2757014
307 | 2766529
308 | 2776045
309 | 2785560
310 | 2795076
311 | 2804591
312 | 2814107
313 | 2823622
314 | 2833137
315 | 2842653
316 | 2852168
317 | 2861684
318 | 2871199
319 | 2880715
320 | 2890230
321 | 2899746
322 | 2909261
323 | 2918777
324 | 2928292
325 | 2937807
326 | 2947323
327 | 2956838
328 | 2966354
329 | 2975869
330 | 2985385
331 | 2994900
332 | 3004416
333 | 3013931
334 | 3023447
335 | 3032962
336 | 3042477
337 | 3051993
338 | 3061508
339 | 3071024
340 | 3080539
341 | 3090055
342 | 3099570
343 | 3109086
344 | 3118601
345 | 3128117
346 | 3137632
347 | 3147147
348 | 3156663
349 | 3166178
350 | 3175694
351 | 3185209
352 | 3194725
353 | 3204240
354 | 3213756
355 | 3223271
356 | 3232787
357 | 3242302
358 | 3251817
359 | 3261333
360 | 3270848
361 | 3280364
362 | 3289879
363 | 3299395
364 | 3308910
365 | 3318426
366 | 3327941
367 | 3337457
368 | 3346972
369 | 3356487
370 | 3366003
371 | 3375518
372 | 3385034
373 | 3394549
374 | 3404065
375 | 3413580
376 | 3423096
377 | 3432611
378 | 3442127
379 | 3451642
380 | 3461157
381 | 3470673
382 | 3480188
383 | 3489704
384 | 3499219
385 | 3508735
386 | 3518250
387 | 3527766
388 | 3537281
389 | 3546797
390 | 3556312
391 | 3565827
392 | 3575343
393 | 3584858
394 | 3594374
395 | 3603889
396 | 3613405
397 | 3622920
398 | 3632436
399 | 3641951
400 | 3651467
401 | 3660982
402 | 3670497
403 | 3680013
404 | 3689528
405 | 3699044
406 | 3708559
407 | 3718075
408 | 3727590
409 | 3737106
410 | 3746621
411 | 3756137
412 | 3765652
413 | 3775167
414 | 3784683
415 | 3794198
416 | 3803714
417 | 3813229
418 | 3822745
419 | 3832260
420 | 3841776
421 | 3851291
422 | 3860807
423 | 3870322
424 | 3879837
425 | 3889353
426 | 3898868
427 | 3908384
428 | 3917899
429 | 3927415
430 | 3936930
431 | 3946446
432 | 3955961
433 | 3965477
434 | 3974992
435 | 3984507
436 | 3994023
437 | 4003538
438 | 4013054
439 | 4022569
440 | 4032085
441 | 4041600
442 | 4051116
443 | 4060631
444 | 4070147
445 | 4079662
446 | 4089157
447 | 4098653
448 | 4108148
449 | 4117643
450 | 4127139
451 | 4136634
452 | 4146129
453 | 4155625
454 | 4165120
455 | 4174615
456 | 4184111
457 | 4193606
458 | 4203101
459 | 4212597
460 | 4222092
461 | 4231587
462 | 4241083
463 | 4250578
464 | 4260073
465 | 4269569
466 | 4279064
467 | 4288559
468 | 4298055
469 | 4307550
470 | 4317045
471 | 4326541
472 | 4336036
473 | 4345531
474 | 4355027
475 | 4364522
476 | 4374018
477 | 4383513
478 | 4393008
479 | 4402504
480 | 4411999
481 | 4421494
482 | 4430990
483 | 4440485
484 | 4449980
485 | 4459476
486 | 4468971
487 | 4478466
488 | 4487962
489 | 4497457
490 | 4506952
491 | 4516448
492 | 4525943
493 | 4535438
494 | 4544934
495 | 4554429
496 | 4563924
497 | 4573420
498 | 4582915
499 | 4592410
500 | 4601906
501 | 4611401
502 | 4620896
503 | 4630392
504 | 4639887
505 | 4649382
506 | 4658878
507 | 4668373
508 | 4677868
509 | 4687364
510 | 4696859
511 | 4706354
512 | 4715850
513 | 4725345
514 | 4734840
515 | 4744336
516 | 4753831
517 | 4763326
518 | 4772822
519 | 4782317
520 | 4791812
521 | 4801308
522 | 4810803
523 | 4820298
524 | 4829794
525 | 4839289
526 | 4848784
527 | 4858280
528 | 4867775
529 | 4877270
530 | 4886766
531 | 4896261
532 | 4905756
533 | 4915252
534 | 4924747
535 | 4934242
536 | 4943738
537 | 4953233
538 | 4962729
539 | 4972224
540 | 4981719
541 | 4991215
542 | 5000710
543 | 5010205
544 | 5019701
545 | 5029196
546 | 5038691
547 | 5048187
548 | 5057682
549 | 5067177
550 | 5076673
551 | 5086168
552 | 5095663
553 | 5105159
554 | 5114654
555 | 5124149
556 | 5133645
557 | 5143140
558 | 5152635
559 | 5162131
560 | 5171626
561 | 5181121
562 | 5190617
563 | 5200112
564 | 5209607
565 | 5219103
566 | 5228598
567 | 5238102
568 | 5247605
569 | 5257109
570 | 5266612
571 | 5276116
572 | 5285619
573 | 5295123
574 | 5304627
575 | 5314130
576 | 5323634
577 | 5333137
578 | 5342641
579 | 5352144
580 | 5361648
581 | 5371152
582 | 5380655
583 | 5390159
584 | 5399662
585 | 5409166
586 | 5418669
587 | 5428173
588 | 5437676
589 | 5447180
590 | 5456684
591 | 5466187
592 | 5475691
593 | 5485194
594 | 5494698
595 | 5504201
596 | 5513705
597 | 5523209
598 | 5532712
599 | 5542216
600 | 5551719
601 | 5561223
602 | 5570726
603 | 5580230
604 | 5589734
605 | 5599237
606 | 5608741
607 | 5618244
608 | 5627748
609 | 5637251
610 | 5646755
611 | 5656259
612 | 5665762
613 | 5675266
614 | 5684769
615 | 5694273
616 | 5703776
617 | 5713280
618 | 5722784
619 | 5732287
620 | 5741791
621 | 5751294
622 | 5760798
623 | 5770301
624 | 5779805
625 | 5789308
626 | 5798812
627 | 5808316
628 | 5817819
629 | 5827323
630 | 5836826
631 | 5846330
632 | 5855833
633 | 5865337
634 | 5874841
635 | 5884344
636 | 5893848
637 | 5903351
638 | 5912855
639 | 5922358
640 | 5931862
641 | 5941366
642 | 5950869
643 | 5960373
644 | 5969876
645 | 5979380
646 | 5988883
647 | 5998387
648 | 6007891
649 | 6017394
650 | 6026898
651 | 6036401
652 | 6045905
653 | 6055408
654 | 6064912
655 | 6074416
656 | 6083919
657 | 6093423
658 | 6102926
659 | 6112430
660 | 6121933
661 | 6131437
662 | 6140940
663 | 6150444
664 | 6159948
665 | 6169451
666 | 6178955
667 | 6188458
668 | 6197962
669 | 6207465
670 | 6216969
671 | 6226473
672 | 6235976
673 | 6245480
674 | 6254983
675 | 6264487
676 | 6273990
677 | 6283494
678 | 6292998
679 | 6302501
680 | 6312005
681 | 6321508
682 | 6331012
683 | 6340515
684 | 6350019
685 | 6359523
686 | 6369026
687 | 6378530
688 | 6388033
689 | 6397537
690 | 6407040
691 | 6416544
692 | 6426048
693 | 6435551
694 | 6445055
695 | 6454558
696 | 6464062
697 | 6473565
698 | 6483069
699 | 6492572
700 | 6502076
701 | 6511580
702 | 6521083
703 | 6530587
704 | 6540090
705 | 6549594
706 | 6559097
707 | 6568601
708 | 6578105
709 | 6587608
710 | 6597112
711 | 6606615
712 | 6616119
713 | 6625622
714 | 6635126
715 | 6644630
716 | 6654133
717 | 6663637
718 | 6673140
719 | 6682644
720 | 6692147
721 | 6701651
722 | 6711155
723 | 6720658
724 | 6730162
725 | 6739665
726 | 6749169
727 | 6758672
728 | 6768176
729 | 6777680
730 | 6787183
731 | 6796687
732 | 6806190
733 | 6815694
734 | 6825197
735 | 6834701
736 | 6844204
737 | 6853708
738 | 6863212
739 | 6872715
740 | 6882219
741 | 6891722
742 | 6901226
743 | 6910729
744 | 6920233
745 | 6929737
746 | 6939240
747 | 6948744
748 | 6958247
749 | 6967751
750 | 6977254
751 | 6986758
752 | 6996266
753 | 7005775
754 | 7015283
755 | 7024791
756 | 7034299
757 | 7043808
758 | 7053316
759 | 7062824
760 | 7072332
761 | 7081841
762 | 7091349
763 | 7100857
764 | 7110365
765 | 7119874
766 | 7129382
767 | 7138890
768 | 7148399
769 | 7157907
770 | 7167415
771 | 7176923
772 | 7186432
773 | 7195940
774 | 7205448
775 | 7214956
776 | 7224465
777 | 7233973
778 | 7243481
779 | 7252989
780 | 7262498
781 | 7272006
782 | 7281514
783 | 7291023
784 | 7300531
785 | 7310039
786 | 7319547
787 | 7329056
788 | 7338564
789 | 7348072
790 | 7357580
791 | 7367089
792 | 7376597
793 | 7386105
794 | 7395614
795 | 7405122
796 | 7414630
797 | 7424138
798 | 7433647
799 | 7443155
800 | 7452663
801 | 7462171
802 | 7471680
803 | 7481188
804 | 7490696
805 | 7500204
806 | 7509713
807 | 7519221
808 | 7528729
809 | 7538238
810 | 7547746
811 | 7557254
812 | 7566762
813 | 7576271
814 | 7585779
815 | 7595287
816 | 7604795
817 | 7614304
818 | 7623812
819 | 7633320
820 | 7642828
821 | 7652337
822 | 7661845
823 | 7671353
824 | 7680862
825 | 7690370
826 | 7699878
827 | 7709386
828 | 7718895
829 | 7728403
830 | 7737911
831 | 7747419
832 | 7756928
833 | 7766436
834 | 7775944
835 | 7785452
836 | 7794961
837 | 7804469
838 | 7813977
839 | 7823486
840 | 7832994
841 | 7842502
842 | 7852010
843 | 7861519
844 | 7871027
845 | 7880535
846 | 7890043
847 | 7899552
848 | 7909060
849 | 7918568
850 | 7928077
851 | 7937585
852 | 7947093
853 | 7956601
854 | 7966110
855 | 7975618
856 | 7985126
857 | 7994634
858 | 8004143
859 | 8013651
860 | 8023159
861 | 8032667
862 | 8042176
863 | 8051684
864 | 8061192
865 | 8070701
866 | 8080209
867 | 8089717
868 | 8099225
869 | 8108734
870 | 8118242
871 | 8127750
872 | 8137258
873 | 8146767
874 | 8156275
875 | 8165783
876 | 8175291
877 | 8184800
878 | 8194308
879 | 8203816
880 | 8213325
881 | 8222833
882 | 8232341
883 | 8241849
884 | 8251358
885 | 8260866
886 | 8270374
887 | 8279882
888 | 8289391
889 | 8298899
890 | 8308407
891 | 8317915
892 | 8327424
893 | 8336932
894 | 8346440
895 | 8355949
896 | 8365457
897 | 8374965
898 | 8384473
899 | 8393982
900 | 8403490
901 | 8412998
902 | 8422506
903 | 8432015
904 | 8441523
905 | 8451031
906 | 8460540
907 | 8470048
908 | 8479556
909 | 8489064
910 | 8498573
911 | 8508081
912 | 8517589
913 | 8527097
914 | 8536606
915 | 8546114
916 | 8555622
917 | 8565130
918 | 8574639
919 | 8584147
920 | 8593655
921 | 8603164
922 | 8612672
923 | 8622180
924 | 8631688
925 | 8641197
926 | 8650705
927 | 8660213
928 | 8669721
929 | 8679230
930 | 8688738
931 | 8698246
932 | 8707754
933 | 8717263
934 | 8726771
935 | 8736291
936 | 8745811
937 | 8755331
938 | 8764851
939 | 8774370
940 | 8783890
941 | 8793410
942 | 8802930
943 | 8812450
944 | 8821970
945 | 8831490
946 | 8841010
947 | 8850530
948 | 8860049
949 | 8869569
950 | 8879089
951 | 8888609
952 | 8898129
953 | 8907649
954 | 8917169
955 | 8926689
956 | 8936209
957 | 8945729
958 | 8955248
959 | 8964768
960 | 8974288
961 | 8983808
962 | 8993328
963 | 9002848
964 | 9012368
965 | 9021888
966 | 9031408
967 | 9040927
968 | 9050447
969 | 9059967
970 | 9069487
971 | 9079007
972 | 9088527
973 | 9098047
974 | 9107567
975 | 9117087
976 | 9126606
977 | 9136126
978 | 9145646
979 | 9155166
980 | 9164686
981 | 9174206
982 | 9183726
983 | 9193246
984 | 9202766
985 | 9212285
986 | 9221805
987 | 9231325
988 | 9240845
989 | 9250365
990 | 9259885
991 | 9269405
992 | 9278925
993 | 9288445
994 | 9297964
995 | 9307484
996 | 9317004
997 | 9326524
998 | 9336044
999 | 9345564
1000 | 9355084
1001 | 9364604
1002 | 9374124
1003 | 9383644
1004 | 9393163
1005 | 9402683
1006 | 9412203
1007 | 9421723
1008 | 9431243
1009 | 9440763
1010 | 9450283
1011 | 9459803
1012 | 9469323
1013 | 9478842
1014 | 9488362
1015 | 9497882
1016 | 9507402
1017 | 9516922
1018 | 9526442
1019 | 9535962
1020 | 9545482
1021 | 9555002
1022 | 9564521
1023 | 9574041
1024 | 9583561
1025 | 9593081
1026 | 9602601
1027 | 9612121
1028 | 9621641
1029 | 9631161
1030 | 9640681
1031 | 9650200
1032 | 9659720
1033 | 9669240
1034 | 9678760
1035 | 9688280
1036 | 9697800
1037 | 9707320
1038 | 9716840
1039 | 9726360
1040 | 9735879
1041 | 9745399
1042 | 9754919
1043 | 9764439
1044 | 9773959
1045 | 9783479
1046 | 9792999
1047 | 9802519
1048 | 9812039
1049 | 9821559
1050 | 9831078
1051 | 9840598
1052 | 9850118
1053 | 9859638
1054 | 9869158
1055 | 9878678
1056 | 9888198
1057 | 9897718
1058 | 9907238
1059 | 9916757
1060 | 9926277
1061 | 9935797
1062 | 9945317
1063 | 9954837
1064 | 9964357
1065 | 9973877
1066 | 9983397
1067 | 9992917
1068 | 10002436
1069 | 10011956
1070 | 10021476
1071 | 10030996
1072 | 10040516
1073 | 10050036
1074 | 10059556
1075 | 10069076
1076 | 10078596
1077 | 10088115
1078 | 10097635
1079 | 10107155
1080 | 10116675
1081 | 10126195
1082 | 10135715
1083 | 10145235
1084 | 10154755
1085 | 10164275
1086 | 10173794
1087 | 10183314
1088 | 10192834
1089 | 10202354
1090 | 10211874
1091 | 10221394
1092 | 10230914
1093 | 10240434
1094 | 10249954
1095 | 10259474
1096 | 10268993
1097 | 10278513
1098 | 10288033
1099 | 10297553
1100 | 10307073
1101 | 10316593
1102 | 10326113
1103 | 10335633
1104 | 10345153
1105 | 10354672
1106 | 10364192
1107 | 10373712
1108 | 10383232
1109 | 10392752
1110 | 10402272
1111 | 10411792
1112 | 10421312
1113 | 10430832
1114 | 10440351
1115 | 10449871
1116 | 10459391
1117 | 10468911
1118 | 10478431
1119 | 10487958
1120 | 10497484
1121 | 10507011
1122 | 10516537
1123 | 10526064
1124 | 10535590
1125 | 10545117
1126 | 10554643
1127 | 10564170
1128 | 10573696
1129 | 10583223
1130 | 10592749
1131 | 10602276
1132 | 10611802
1133 | 10621329
1134 | 10630855
1135 | 10640382
1136 | 10649908
1137 | 10659435
1138 | 10668961
1139 | 10678488
1140 | 10688014
1141 | 10697541
1142 | 10707067
1143 | 10716594
1144 | 10726120
1145 | 10735647
1146 | 10745173
1147 | 10754700
1148 | 10764226
1149 | 10773753
1150 | 10783279

## Strαpping Frαksi (hidden)
Plant | Unit_kode | Tank_no | Dari | Sampai | Dari | Sampai | Tinggi | Volume
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 1 | 951.3
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 2 | 1902.6
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 3 | 2854
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 4 | 3805.3
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 5 | 4756.6
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 6 | 5707.9
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 7 | 6659.2
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 8 | 7610.6
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 9 | 8561.9
R501 | 1 | T-902 | 0.5 | 1.68 | 50 | 168 | 10 | 9513.2
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 1 | 949.3
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 2 | 1898.6
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 3 | 2847.9
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 4 | 3797.2
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 5 | 4746.5
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 6 | 5695.8
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 7 | 6645.1
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 8 | 7594.4
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 9 | 8543.7
R501 | 1 | T-902 | 1.68 | 2.91 | 168 | 291 | 10 | 9493
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 1 | 951.5
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 2 | 1903.1
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 3 | 2854.6
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 4 | 3806.2
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 5 | 4757.7
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 6 | 5709.3
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 7 | 6660.8
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 8 | 7612.4
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 9 | 8563.9
R501 | 1 | T-902 | 2.91 | 4.45 | 291 | 445 | 10 | 9515.5
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 1 | 949.5
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 2 | 1899.1
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 3 | 2848.6
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 4 | 3798.1
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 5 | 4747.7
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 6 | 5697.2
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 7 | 6646.7
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 8 | 7596.3
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 9 | 8545.8
R501 | 1 | T-902 | 4.45 | 5.66 | 445 | 566 | 10 | 9495.3
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 1 | 950.4
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 2 | 1900.7
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 3 | 2851.1
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 4 | 3801.4
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 5 | 4751.8
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 6 | 5702.1
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 7 | 6652.5
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 8 | 7602.9
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 9 | 8553.2
R501 | 1 | T-902 | 5.66 | 7.51 | 566 | 751 | 10 | 9503.6
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 1 | 950.8
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 2 | 1901.7
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 3 | 2852.5
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 4 | 3803.3
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 5 | 4754.1
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 6 | 5705
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 7 | 6655.8
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 8 | 7606.6
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 9 | 8557.4
R501 | 1 | T-902 | 7.51 | 9.34 | 751 | 934 | 10 | 9508.3
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 1 | 952
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 2 | 1904
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 3 | 2856
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 4 | 3808
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 5 | 4759.9
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 6 | 5711.9
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 7 | 6663.9
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 8 | 7615.9
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 9 | 8567.9
R501 | 1 | T-902 | 9.34 | 11.18 | 934 | 1118 | 10 | 9519.9
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 1 | 952.7
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 2 | 1903.3
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 3 | 2858
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 4 | 3810.6
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 5 | 4763.3
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 6 | 5715.9
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 7 | 6668.6
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 8 | 7621.2
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 9 | 8573.9
R501 | 1 | T-902 | 11.18 | 11.5 | 1118 | 1150 | 10 | 9526.5
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
[8 empty] | 0 | D
0
0
0
0
0
0
0
0
0
[8 empty] | 0 | H
0
0
0
0
0
0
0
0
0
[8 empty] | 0 | L
0
0
0
0
0
0
0
0
0
[8 empty] | 0 | P
0
0
0
0
0
0
0
0
0
[8 empty] | 0 | T
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Pekerja (hidden)
DAFTAR TANDA TERIMA
PENGGANTI UANG MAKAN PHISYCAL CHEK TRIWULAN KE-4
No | Nopek | Nama | Bagian | Jabatan
| 749767 | Cipta Ariandi Lubis | Asset Management | Jr Officer Land Affairs RU V
| 748827 | Eko Dewanto | Asset Management | Officer Asset Optimization RU V
| 672287 | Pulo Rencus Manurung | Asset Management | Asset Management RU V Area Manager
| 746815 | Puspa Dewi Shinta Rukmi | Asset Management | Officer Land Affairs RU V
| 749613 | Adhya Husni Mubarak | BP Refinery | Junior Officer BP Refinery
| 751681 | Ninda Sasmita Putri | BP Refinery | Junior Officer BP Refinery
| 651097 | Nur Hasanah | BP Refinery | Officer BP Refinery
| 746771 | Nurdin Tri Wibowo | BP Refinery | Senior Officer BP Refinery
| 748634 | Chrisavitri Yulihastiwi | Budget & Performance | Jr Analyst Hydrocarbon Reconsiliation
| 747307 | Dwi Nur Rohmadhoni | Budget & Performance | Jr Assistant Budget Reporting
| 746456 | Endro Alfiansyah | Budget & Performance | Jr Assistant Data & Information System
| 610831 | Irianto | Budget & Performance | Sr Supervisor Refinery Performance
| 720327 | Nasruddin | Budget & Performance | Section Head Budget & Performance
| 646261 | Yulidar Suryanata | Budget & Performance | Sr Supervisor Budget & Controlling
| 752099 | Marselina Tando | Bus Ops & Tech | Assistant Business Support Operations
| 749095 | Muhammad Royamin | Bus Ops & Tech | Jr Assistant Comp Dev & Creative Tech
| 749590 | Prima Kurniawan | Bus Ops & Tech | Assistant Data Center Ops & Automation
| 749848 | Dzikri Fadhilah | Bus Supp & Infra Tech | Jr Assistant Fixed & Mobile Comm Ops
| 748005 | I Komang Trisna Jayadi | Bus Supp & Infra Tech | Officer Network & System Support
| 745718 | Mecky Firmansyah | Bus Supp & Infra Tech | Assistant Fixed & Mobile Comm Ops
| 651997 | Suwarno M. | Bus Supp & Infra Tech | Ast Mgr Bus Supp & Infrastructure Tech
| 751687 | Bagja Mahendra | Comm. & Relation | Jr Officer Comm. & Relations Kalimantan
| 744607 | Dian Hapsari Firasati | Comm. & Relation | Area Manager Comm.& Relations Kalimantan
| 749458 | Imam Rismanto | Comm. & Relation | Head of Comm. & Relations RU V
| 751690 | Keishkara Hanandhita Putri | Comm. & Relation | Jr Officer Comm. & Relations RU V
| 10015567 | Kunadi | Contract Office | Sr Supervisor Bidding
| 752572 | Novia Wulandari | Contract Office | Jr Officer Bidding Maintenance
| 713597 | Ruddy Teguh Prakoso | Contract Office | Section Head Contract Office
| 750012 | Ruri Febriana Sary | Contract Office | Jr Assistant Vendor Master & Adm Ctrl.
| 10017722 | Soeyadi | Contract Office | Sr Supervisor Contract
| 650343 | Tuti Wahyuni | Contract Office | Jr Officer Contract Operational Services
| 674977 | Wiwiek Sulastri | Contract Office | Jr Officer Contract TA & OH Services
| 10017737 | Zulkifli Ali Saat | Contract Office | Jr Officer Bidding TA & OH
| 688909 | A.S.M.Fahrial | Dis. & Wax. | Sr Operator II Tank Dis Wax & Data
| 746284 | Achmad Fihris | Dis. & Wax. | Jr Operator II CDU/HVU
| 746327 | Addi Supriyadi | Dis. & Wax. | Jr Operator II CDU/HVU
| 752823 | Adi Kuswoyo | Dis. & Wax. | Jr Operator I CDU/HVU
| 737549 | Adrian Rachmansyah | Dis. & Wax. | Panelman I CDU/HVU
| 731351 | Ahmad Mutaqin | Dis. & Wax. | Officer Facility & Quality
| 10015213 | Akhmad Siswanto | Dis. & Wax. | Shift Supervisor CDU/HVU
| 624512 | Ardiansyah M. | Dis. & Wax. | Shift Supervisor CDU/HVU
| 695453 | Asnan | Dis. & Wax. | Sr Operator II Site EWTP
| 750276 | Bahri Fahrizal | Dis. & Wax. | Operator I CDU/HVU
| 10015002 | Budiyanto | Dis. & Wax. | Shift Supervisor Wax
| 695283 | Chairuddin | Dis. & Wax. | Panelman II Wax
| 695364 | Dang Marta | Dis. & Wax. | Panelman II CDU/HVU
| 746300 | Diaz Aulivyan | Dis. & Wax. | Jr Operator II CDU/HVU
| 10015218 | Djaelani | Dis. & Wax. | Shift Supervisor Wax
| 745251 | Endik Purnomo | Dis. & Wax. | Operator I CDU/HVU
| 751460 | Faruq Rahmatullah | Dis. & Wax. | Jr Operator I CDU/HVU
| 751462 | Fernandus Fendy Rianto | Dis. & Wax. | Jr Operator I CDU/HVU
| 751114 | Firman Nusantara | Dis. & Wax. | Jr Operator I Site DHP
| 668537 | Gatot Sutrisno | Dis. & Wax. | Panelman II CDU/HVU
| 746326 | Ghafur Rachmadani | Dis. & Wax. | Jr Operator II CDU/HVU
| 10015004 | Gustaf H.S. | Dis. & Wax. | Sr Supervisor EWTP/DHP
| 751112 | Hariyanto | Dis. & Wax. | Jr Operator I CDU/HVU
| 714625 | Herman Kwandang | Dis. & Wax. | Section Head Dis & Wax
| 751464 | Ikrar Tri Putra | Dis. & Wax. | Jr Operator I CDU/HVU
| 751465 | Jaya Murti Widyardi | Dis. & Wax. | Jr Operator I CDU/HVU
| 10015749 | Jemmy Berty Kusen | Dis. & Wax. | Shift Supervisor EWTP/DHP
| 731968 | Jonson Anekke Hutajulu | Dis. & Wax. | Panelman II Wax
| 753387 | Lukman Hakim | Dis. & Wax. | Operator I Site EWTP
| 752809 | Lutfi | Dis. & Wax. | Jr Operator I CDU/HVU
| 747091 | M. Fahrun Richardsa L. | Dis. & Wax. | Operator II Wax
| 731992 | M.W.Henri Butarbutar | Dis. & Wax. | Panelman II CDU/HVU
| 746324 | Markus Kelvin H. Tampubolon | Dis. & Wax. | Jr Operator II CDU/HVU
| 749995 | Masruri Azis | Dis. & Wax. | Operator I CDU/HVU
| 747094 | Maulana Taufik | Dis. & Wax. | Operator II Wax
| 745248 | Meirizal Fahmi | Dis. & Wax. | Operator I CDU/HVU
| 751468 | Moch Farizal Rhomawan | Dis. & Wax. | Jr Operator I CDU/HVU
| 10016298 | Muh Taufik Kohor | Dis. & Wax. | Panelman II Wax
| 751470 | Muhamad Alif Adlu Haq | Dis. & Wax. | Jr Operator I CDU/HVU
| 746328 | Muhammad Nur Iedi | Dis. & Wax. | Jr Operator II CDU/HVU
| 750331 | Muhammad Nurman Febrian | Dis. & Wax. | Operator I CDU/HVU
| 752674 | Muhammad Suryo Hadi Rifai | Dis. & Wax. | Jr Operator I Site EWTP
| 746291 | Muhammad Taufik Arrachman | Dis. & Wax. | Jr Operator II CDU/HVU
| 695494 | Novizon | Dis. & Wax. | Panelman II Wax
| 752831 | Pandu Aditya W.U. | Dis. & Wax. | Jr Operator I CDU/HVU
| 624829 | Perngadi | Dis. & Wax. | Officer Facility & Quality
| 10018062 | Petrus Sanda S | Dis. & Wax. | Sr Operator II Site EWTP
| 687556 | Rahardjo | Dis. & Wax. | Panelman II Wax
| 748335 | Rahman Effendi | Dis. & Wax. | Panelman I CDU/HVU
| 752811 | Rendra Windyanto | Dis. & Wax. | Jr Operator I Wax
| 679286 | Ruslan | Dis. & Wax. | Panelman II Wax
| 10018045 | Saparudin | Dis. & Wax. | Shift Supervisor EWTP/DHP
| 746248 | Septian Trinanda | Dis. & Wax. | Jr Operator II CDU/HVU
| 637254 | Sobrah Kuntjoro | Dis. & Wax. | Sr Operator II Site EWTP
| 610978 | Sudjito | Dis. & Wax. | Shift Supervisor EWTP/DHP
| 688893 | Sugianto | Dis. & Wax. | Sr Operator II Site DHP
| 625071 | Suharto | Dis. & Wax. | Sr Supervisor Process Dis & Wax
| 10016103 | Sukisno | Dis. & Wax. | Shift Supervisor EWTP/DHP
| 743867 | Su'udi Ariyanto | Dis. & Wax. | Operator II CDU/HVU
| 10016820 | Suwanto | Dis. & Wax. | Shift Supervisor CDU/HVU
| 10016101 | Syamsuri | Dis. & Wax. | Shift Supervisor CDU/HVU
| 625306 | Ulim Maful Priyanto | Dis. & Wax. | Shift Supervisor Wax
| 731205 | Vordy Hence Ropa | Dis. & Wax. | Panelman II CDU/HVU
| 751486 | Yahya Sahru Ramadhan | Dis. & Wax. | Jr Operator I CDU/HVU
| 731173 | Zaenal Abidin | Dis. & Wax. | Panelman II Wax
| 750121 | Agung Budi P | Dockyard | Jr. Spv. Adm Dockyard
| 642681 | Abdul Manaf | Elec. & Inst. Insp.Engineer | Engineer Instrument Inspection
| 750536 | Aditya Rifki Dwi Pamungkas Pur | Elec. & Inst. Insp.Engineer | Jr Engineer I Instrument Inspection
| 753242 | Afiat Dirghantara | Elec. & Inst. Insp.Engineer | Jr Engineer I Electrical Inspection
| 751802 | Antonius Allan Darmasaputra | Elec. & Inst. Insp.Engineer | Jr Engineer I Instrument Inspection
| 752025 | Aryo Wibisono | Elec. & Inst. Insp.Engineer | Jr Engineer I Electrical Inspection
| 747976 | Audithira Mahardriasmoko | Elec. & Inst. Insp.Engineer | Group Leader Instrument Inspection
| 753249 | Daniar Rizki Wijayanti | Elec. & Inst. Insp.Engineer | Jr Engineer I Electrical Inspection
| 753252 | Mirza Rahim | Elec. & Inst. Insp.Engineer | Jr Engineer I Electrical Inspection
| 749830 | Mochamad Nurdiansyah | Elec. & Inst. Insp.Engineer | Jr Engineer II Instrument Inspection
| 725569 | Nanang Supriatna | Elec. & Inst. Insp.Engineer | Engineer Instrument Inspection
| 749838 | Randy Saut Purba | Elec. & Inst. Insp.Engineer | Jr Engineer II Instrument Inspection
| 608475 | Ronny Kumendong | Elec. & Inst. Insp.Engineer | Group Leader Electrical Inspection
| 719997 | Sumiatno | Elec. & Inst. Insp.Engineer | Lead of Electrical & Instrument Engineer
| 751839 | Suwega Drestantiarto | Elec. & Inst. Insp.Engineer | Jr Engineer I Instrument Inspection
| 608264 | Syahmad | Elec. & Inst. Insp.Engineer | Engineer Electrical Inspection
| 642179 | Untung Samiadji | Elec. & Inst. Insp.Engineer | Engineer Electrical Inspection
| 750574 | Yoga Budi Setiawan | Elec. & Inst. Insp.Engineer | Jr Engineer I Instrument Inspection
| 752190 | Yusup Suryadi | Elec. & Inst. Insp.Engineer | Jr Engineer I Electrical Inspection
| 714593 | Arif Budiyono | Energy Cons. Loss Ctr. | Section Head Energy Con & Loss Control
| 679553 | Rizal Alriansyah | Energy Cons. Loss Ctr. | Sr Supervisor Energy & Combustion
| 740026 | Edy Januari Utama | Eng. & Dev. | Manager Engineering & Development
| 608701 | Herman Sudiono | Eng. & Dev. | Sr Engineer Reforming & Treating
| 746927 | Iman Sukmajati | Eng. & Dev. | Sr Engineer CDU, UTL & Offsite
| 10017494 | Micky Paulus Walangare | Eng. & Dev. | Sr Engineer HCU & H2 Plant
| 747803 | Dessy Chaterina Ullyma Gultom | Environment | Sr Supervisor Monitoring Waste Mgt.
| 641944 | Djoko Soenaryo | Environment | Sr Supervisor Environ. Mon. & Evaluation
| 714522 | Fery Purnawantoko | Environment | Section Head Environment
| 639547 | Hanafiah | Environment | Jr Officer Environmental Monitoring
| 747566 | Ika Yuliastuti | Environment | Jr Officer Environmental Dev. & Reg.
| 604457 | P M Rahayu | Environment | Jr Officer Environmental Evaluation
| 731724 | Rudi Setyawan | Environment | Jr Officer Hazardous Waste Control
| 679189 | Said Abdullah | Environment | Jr Officer Non Hazardous Waste Control
| 10017684 | Hendy Suhendi | Equipment Overhaul | Engineer Mechanical O/H Planning
| 694919 | Tahal Pasni | Equipment Overhaul | Engineer II RE O/H Planning
| 714674 | Yuliwis | Equipment Overhaul | Section Head Equipment Overhaul
| 679545 | Djoko Permadi | Equipment Reliability | Engineer I Instr. Equip. Reliability
| 636614 | Joko Sinaga | Equipment Reliability | Engineer I Rotating Equip Reliabilty
| 679472 | Mugiyono | Equipment Reliability | Engineer I Elect.Equip. Reliability
| 625493 | Muhammad Nasrullah | Equipment Reliability | Engineer II Instr. Equip. Reliability
| 641952 | Muhammad Yusuf | Equipment Reliability | Engineer I Stat. Equip. Reliability
| 607981 | Sugeng | Equipment Reliability | Engineer II Rotating Equip Reliabilty
| 718343 | Supriyono | Equipment Reliability | Section Head Equipment Reliability
| 653479 | Tri Bambang H | Equipment Reliability | Engineer II Elect.Equip. Reliability
| 745088 | Agustina Kusumaningtyas | Finance Business Support | Sr Spv Oil Custody&Admin Clearance
| 733433 | Arief Maukar R. | Finance Business Support | Sr Supervisor Intake Output
| 752384 | Aulia Rahman | Finance Business Support | Jr Assistant Shipping Shift
| 677763 | Haba'udin | Finance Business Support | Ast Manager Finance Business Supp.
| 750410 | Harlan Setiadi | Finance Business Support | JrOfficer MtrlBlnce,OilDiscre&OilMovmnt
| 750413 | Lovasta Holy Marciano | Finance Business Support | Jr Officer Expenditure
| 749306 | Mega Rahmasari | Finance Business Support | Analyst Consolidation & Reporting
| 752386 | Muhammad Romli | Finance Business Support | Jr Assistant Shipping Shift
| 750418 | Pralabdha Ananggadipa | Finance Business Support | Analyst RU Fin Perf,ProductCosting&Invtr
| 752388 | Wendy Wewisa P | Finance Business Support | Jr Assistant Shipping Shift
| 751631 | Erwin Tri Hutomo | Finance Refinery Sorong | Jr Officer Budget & Financial Performance
| 701357 | Kamson Hutahayan | Finance Refinery Sorong | Section Head Finance Refinery Sorong
| 750638 | Muhammad Ihsan Riyadi | Finance Refinery Sorong | Jr Officer Cash, Bank & L/C
| 751640 | Oktavianus Aditya Mahendra | Finance Refinery Sorong | Jr Analyst Tax Accounting
| 745124 | Widy Nugroho | Finance Refinery Sorong | Jr Officer AP-AR
| 728947 | Edi Budiarto | Finance RU V | Finance RU V Manager
| 100018605 | Aditya H.R. | Financial Acc. & Treasury
| 732056 | Chananoen | Financial Acc. & Treasury | Officer Tax Audit & Advisory
| 745123 | Daniyandris Idris | Financial Acc. & Treasury | Jr Officer Cash, Bank & L/C
| 711344 | Edy Fatman | Financial Acc. & Treasury | Ast Manager Financial Acc. & Treasury
| 750408 | Ganjar Nugraha | Financial Acc. & Treasury | Jr Officer Asset Acc, Mterial&Dtabse Mgt
| 731165 | Hadi Sutrisno | Financial Acc. & Treasury | Sr Supervisor General Accounting
| 745119 | Ika Sri Wahyuni | Financial Acc. & Treasury | Jr Officer AP
| 745116 | Muhammad Zulhaidian Nor | Financial Acc. & Treasury | Jr Officer AP
| 657294 | Nilawati | Financial Acc. & Treasury | Sr Supervisor AP - AR
| 752161 | Susetyo Adhy Nugroho | Financial Acc. & Treasury | Jr Officer Employee AP-AR & Non Trade AR
| 752843 | Achmat Hanafi Mawanda | Fire and Insurance | Jr Fireman I RU V - Balikpapan
| 752754 | Ade Krisna Erlangga | Fire and Insurance | Jr Fireman I RU V - Lawe2
| 746307 | Adhika Noegroho | Fire and Insurance | Jr Fireman II RU V - Balikpapan
| 752895 | Arief Nurhalid | Fire and Insurance | Jr Fireman I RU V - Balikpapan
| 748348 | Aulia Rachim | Fire and Insurance | Fireman II RU V - Balikpapan
| 746461 | Boy Bayu Adi | Fire and Insurance | Fireman II RU V - Balikpapan
| 746296 | Budi Susanto | Fire and Insurance | Jr Fireman II RU V - Balikpapan
| 748346 | Danding Adijaya | Fire and Insurance | Fireman II RU V - Balikpapan
| 748349 | Dean Werena | Fire and Insurance | Fireman II RU V - Balikpapan
| 731895 | Djaenuri | Fire and Insurance | Jr Officer Fire Facilities & Equipment
| 752802 | Eko Priyadi | Fire and Insurance | Jr Fireman I Lawe2
| 748345 | Erliza Achmad Akbar | Fire and Insurance | Fireman II RU V - Balikpapan
| 748350 | Hadi Yulianto | Fire and Insurance | Fireman II RU V - Lawe2
| 746278 | Iyus Setiawan Suherman | Fire and Insurance | Jr Fireman II RU V - Balikpapan
| 10015568 | Jumadan | Fire and Insurance | Sr Fireman I RU V - Lawe2
| 748347 | Muhammad Chakim Afrizal | Fire and Insurance | Fireman II RU V - Balikpapan
| 746312 | Murip | Fire and Insurance | Jr Fireman II RU V - Lawe2
| 713831 | Nurmatias | Fire and Insurance | Section Head Fire & Insurance
| 752820 | Ogi Hendrawan | Fire and Insurance | Jr Fireman I RU V - Lawe2
| 748709 | Oka Adhitya Kusumawardhana | Fire and Insurance | Jr Officer Insurance
| 752901 | Panji Sanjaya | Fire and Insurance | Jr Fireman I RU V - Balikpapan
| 748351 | Putut Widiyanto | Fire and Insurance | Fireman II RU V - Balikpapan
| 746459 | Setyo Widodo | Fire and Insurance | Fireman II RU V - Balikpapan
| 750360 | Simion Ariadno Sitompul | Fire and Insurance | Jr Assistant Training & Administration
| 725155 | Sri Martono | Fire and Insurance | Sr Supervisor Insurance & Fire Facility
| 10015745 | Sukardi | Fire and Insurance | Sr Fireman II RU V - Lawe2
| 752673 | Surya Agung Wirawan | Fire and Insurance | Jr Fireman I RU V - Balikpapan
| 752839 | Surya Arief Bayu Wijaya | Fire and Insurance | Jr Fireman I RU V - Balikpapan
| 719842 | Eko Hernanto | General Affairs | Manager General Affairs
| 674596 | Agus Y. Parlan | General Maint. | Section Head General Maintenance
| 10016104 | Asrani | General Maint. | Sr Technician II Civil Building
| 687515 | Dwi Edhi Purwanto | General Maint. | Sr Technician II Civil Facility & Const
| 636809 | Hamdan | General Maint. | Sr Technician II Civil Facility & Const
| 737516 | Ismed Agussani | General Maint. | Sr Technician II Civil Building
| 10017678 | Jumansyah | General Maint. | Sr Supervisor Civil & Const
| 10017723 | Syachran Sayid Andi A.B | General Maint. | Sr Supervisor Elect & Inst
| 10017010 | Zaenal Abidin | General Maint. | Sr Technician II Instrument
| 10017747 | Fawziah Mulia Savitri | General Manager | Staf Utama Diperbantukan
| 709093 | Yulian Dekri | General Manager | GM Refinery Unit V - Balikpapan
| 750248 | Adi Arianto Ramadani Keni | HCC | Operator I Fractionator HCU
| 10016817 | Aditiawarman | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 748528 | Aditya Pradikta Novianto | HCC | Panelman I Reactor HCU & Common
| 737484 | Agus Hartono | HCC | Sr Operator II HPU - 8ABC/H2Rec/SWS
| 10015001 | Agus Jarwanto | HCC | Panelman II Fractionator HCU
| 10018059 | Agus Praptono | HCC | Sr Operator II Reactor HCU & Common
| 747093 | Akbar | HCC | Operator II Fractionator HCU
| 10016641 | Alla Widiarto | HCC | Shift Supervisor HPU - 8ABC/H2Rec/SWS
| 737532 | Aloysahid Korompis D. | HCC | Panelman II Reactor HCU & Common
| 10014999 | Aminul Harahap | HCC | Panelman II HVU II Comfac & FGHRS
| 751443 | Anas Akbar | HCC | Jr Operator I Fractionator HCU
| 747304 | Antony Lumban Gaol | HCC | Operator II HVU II Comfac & FGHRS
| 751447 | Anugrah | HCC | Jr Operator I Fractionator HCU
| 748527 | Ardhi Kartika Yoga | HCC | Panelman I Reactor HCU & Common
| 625209 | Ardiansyah | HCC | Shift Supervisor Reactor HCU & Common
| 10016815 | Armadan | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 751449 | Azkha Dezthara | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 679829 | Bardiansyah | HCC | Shift Supervisor Fractionator HCU
| 745257 | Bastarie | HCC | Jr Operator II HPU - 8ABC/H2Rec/SWS
| 751451 | Bayhaqi Fauzi | HCC | Jr Operator I HVU II Comfac & FGHRS
| 751108 | Bayu Umbara | HCC | Jr Operator I HVU II Comfac & FGHRS
| 751452 | Benhard Maruli Butar Butar | HCC | Jr Operator I Fractionator HCU
| 752903 | Bismo Ilham Saputro | HCC | Jr Operator I Fractionator HCU
| 746454 | Choirul Anam | HCC | Operator II HPU - 8ABC/H2Rec/SWS
| 753414 | Chrisna Yudha Putra | HCC | Operator I Reactor HCU & Common
| 751455 | Dana Satrian | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 753513 | Deden Riskianandika | HCC | Operator I Reactor HCU & Common
| 732023 | Delmi Wardi | HCC | Sr Operator II HPU - 8ABC/H2Rec/SWS
| 750292 | Dimas Adhitya Rahman | HCC | Operator I Fractionator HCU
| 743859 | Eddy Budiyanto | HCC | Operator II Reactor HCU & Common
| 636493 | Eddy Setyono | HCC | Shift Supervisor HPU - 8ABC/H2Rec/SWS
| 745253 | Edi Putrawan | HCC | Jr Operator II HVU II Comfac & FGHRS
| 679659 | Edy Suprapto | HCC | Panelman II Fractionator HCU
| 10015216 | Edy Suryono | HCC | Sr Operator II HVU II Comfac & FGHRS
| 745258 | Eko Mulyanto | HCC | Operator I HPU - 8ABC/H2Rec/SWS
| 752840 | Eko Pradito | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 679391 | Elpon Ritonga | HCC | Panelman II Fractionator HCU
| 752842 | Erick Ananta Wijaya | HCC | Jr Operator I Fractionator HCU
| 679926 | F. Nainggolan | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 748343 | Fachruddin Syahrul | HCC | Operator II HVU II Comfac & FGHRS
| 752750 | Fahd Syah Alam | HCC | Jr Operator I Fractionator HCU
| 746455 | Fahrul Kuswanto | HCC | Operator II HVU II Comfac & FGHRS
| 745247 | Fandy Hidayat | HCC | Operator I HPU - 8ABC/H2Rec/SWS
| 751459 | Faris Hafishuddin | HCC | Jr Operator I HVU II Comfac & FGHRS
| 745255 | Febro | HCC | Jr Operator II Reactor HCU & Common
| 642187 | Firmananto W.W. | HCC | Shift Supervisor HVU II, Comfac & FGHRS
| 752845 | Frans Adiyoga Kurniawan | HCC | Jr Operator I HVU II Comfac & FGHRS
| 749552 | Fromesa Wisnu Tri Pamungkas | HCC | Operator II Reactor HCU & Common
| 752889 | Givari Anggariyanto | HCC | Jr Operator I HVU II Comfac & FGHRS
| 753408 | Gusti Arsy A.J Bilfani | HCC | Operator I Reactor HCU & Common
| 732015 | Gustiawarman | HCC | Panelman II Reactor HCU & Common
| 747794 | Hamka Mustika Muhayar | HCC | Operator II Reactor HCU & Common
| 750307 | Hamsah | HCC | Operator I HPU - 8ABC/H2Rec/SWS
| 679537 | Hapriansyah | HCC | Sr Operator II HVU II Comfac & FGHRS
| 636688 | Haryanto | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 636711 | Herling E Sumolang | HCC | Shift Supervisor Fractionator HCU
| 743786 | Hermanto | HCC | Operator II HVU II Comfac & FGHRS
| 679797 | Husaini | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 752800 | Ibnu Heriansyah | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 10015569 | Idrus Mulyono | HCC | Panelman II Reactor HCU & Common
| 752753 | Iliyyanda Fahrian | HCC | Jr Operator I HVU II Comfac & FGHRS
| 745256 | Iwan Darmawan | HCC | Jr Operator II Fractionator HCU
| 10016304 | Jayani | HCC | Shift Supervisor HVU II, Comfac & FGHRS
| 745261 | Joko Sutrisno | HCC | Operator I Reactor HCU & Common
| 624772 | Judo | HCC | Panelman II Fractionator HCU
| 746457 | Kaspul Anwar | HCC | Operator II Reactor HCU & Common
| 751467 | Khamarudin | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 752670 | Luky Aviansyah Nur | HCC | Jr Operator I HVU II Comfac & FGHRS
| 752841 | M. Yusni Ari Setiawan | HCC | Jr Operator I Fractionator HCU
| 679845 | Mardi | HCC | Sr Operator II HPU - 8ABC/H2Rec/SWS
| 731198 | Mardi Basuki | HCC | Panelman II Reactor HCU & Common
| 747095 | Mochamad Ferdy Abdillah | HCC | Operator II HVU II Comfac & FGHRS
| 637246 | Moh Mashud | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 679464 | Moh.Japri | HCC | Panelman II Fractionator HCU
| 679731 | Mugianto | HCC | Panelman II HVU II Comfac & FGHRS
| 731295 | Mugiyanto | HCC | Panelman II HVU II Comfac & FGHRS
| 748337 | Muhammad Rahmadani | HCC | Operator II Reactor HCU & Common
| 636622 | Mursidi | HCC | Sr Supervisor HVU II/Comfac & FGHRS
| 679626 | Nanang P.S Wahyudi | HCC | Panelman II Fractionator HCU
| 752806 | Nataliansyah Malik | HCC | Jr Operator I Fractionator HCU
| 679691 | Norhidayat M. | HCC | Sr Operator II Reactor HCU & Common
| 743794 | Noviansyah | HCC | Operator II HVU II Comfac & FGHRS
| 748529 | Nuclies Aquanto | HCC | Panelman I HPU - 8ABC/H2Rec/SWS
| 637068 | Nuryanto | HCC | Panelman II Fractionator HCU
| 725577 | Pepen Permana | HCC | Panelman II HVU II Comfac & FGHRS
| 639944 | Prio Anggodo | HCC | Shift Supervisor HPU - 8ABC/H2Rec/SWS
| 10015219 | Pudjo Suseno | HCC | Shift Supervisor HVU II, Comfac & FGHRS
| 679212 | Puspa Wardaya | HCC | Sr Operator II HPU - 8ABC/H2Rec/SWS
| 731376 | Ramelan Beny Prastyanto | HCC | Panelman II HVU II Comfac & FGHRS
| 731181 | Resmihadi | HCC | Officer Facility & Quality
| 747305 | Riduan Simanjuntak | HCC | Operator II Reactor HCU & Common
| 747798 | Ridwan Abdus | HCC | Operator II HVU II Comfac & FGHRS
| 642462 | Riswandri | HCC | Shift Supervisor Fractionator HCU
| 10018620 | Rochamad Subandi | HCC | Sr Supervisor HPU - 8ABC/H2Rec/SWS
| 637043 | Rostan | HCC | Shift Supervisor HPU - 8ABC/H2Rec/SWS
| 637084 | Rudianto | HCC | Shift Supervisor Reactor HCU & Common
| 751482 | Sandi Pradana | HCC | Jr Operator I Fractionator HCU
| 741314 | Setyo Pitoyo | HCC | Section Head HydroCracking Complex
| 679415 | Simbolon H. | HCC | Panelman II HPU - 8ABC/H2Rec/SWS
| 745246 | Sudarmono | HCC | Jr Operator II HPU - 8ABC/H2Rec/SWS
| 637076 | Sudarno | HCC | Sr Supervisor Process HC Unibon
| 636574 | Sudarwanto | HCC | Shift Supervisor Fractionator HCU
| 695389 | Suharsono | HCC | Sr Operator II Reactor HCU & Common
| 731384 | Sumadi | HCC | Panelman II Reactor HCU & Common
| 10015746 | Sunaryo | HCC | Shift Supervisor HVU II, Comfac & FGHRS
| 695478 | Supardiyono | HCC | Panelman II HVU II Comfac & FGHRS
| 751109 | Suprapto | HCC | Jr Operator I Fractionator HCU
| 636639 | Suprayogi | HCC | Shift Supervisor Reactor HCU & Common
| 752244 | Supriyanto | HCC | Jr Operator I Reactor HCU & Common
| 752829 | Syarifuddin Saenong | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 751484 | Thomi Irfanmulia | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 743761 | Tris Fitrianto | HCC | Operator II Fractionator HCU
| 748339 | Vivit Kristian Saputra | HCC | Operator II HPU - 8ABC/H2Rec/SWS
| 751485 | Wahyu Nugraha | HCC | Jr Operator I Fractionator HCU
| 747797 | Wendy Eka Wardani | HCC | Operator II HPU - 8ABC/H2Rec/SWS
| 747495 | Yesay Setiawan | HCC | Officer Facility & Quality
| 751487 | Yoga Pratama | HCC | Jr Operator I HVU II Comfac & FGHRS
| 751488 | Yonda Canggih Atmajanu | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 753516 | Yuda Nugraha | HCC | Operator I Reactor HCU & Common
| 751113 | Zainal Arifin | HCC | Jr Operator I HPU - 8ABC/H2Rec/SWS
| 701884 | Nelson S Sihombing | HR Operation | Senior Supervisor HR Operation
| 752585 | Redy Erdiyanto | HR Operation | Junior Assistant HR Services
| 729538 | Supendi | HR Operation | Assistant HR Services
| 753373 | Adrianto Bagus Sulaksono | HSC | Operator I Distill. & Separation
| 645557 | Agus Bekti Wibowo | HSC | Panelman II Distill. & Separation
| 751440 | Ahmad Iskandar | HSC | Jr Operator I Reforming
| 751441 | Albert Christian Silaban | HSC | Jr Operator I Reforming
| 668489 | Amboha | HSC | Sr Operator II Distill. & Separation
| 731521 | Amirudin | HSC | Panelman II Distill. & Separation
| 746292 | Andika Mulya Pratama | HSC | Jr Operator II Distill. & Separation
| 624472 | Anwar Achmad Kasim | HSC | Shift Supervisor Reforming
| 746460 | Arcino | HSC | Operator II Reforming
| 750271 | Aristya Eka Putra | HSC | Operator I Distill. & Separation
| 10017012 | Ashari | HSC | Shift Supervisor Reforming
| 10016818 | Bambang Sutoto | HSC | Panelman II Distill. & Separation
| 679497 | Burhanudin | HSC | Sr Operator II Distill. & Separation
| 695186 | Dedi Kuswandi | HSC | Panelman II Distill. & Separation
| 752807 | Didik Junaidi | HSC | Jr Operator I Distill. & Separation
| 10017008 | Djamaluddin | HSC | Panelman II Distill. & Separation
| 10018623 | Djarkasih Polii | HSC | Panelman II Reforming
| 731392 | Djoko Widiyarto | HSC | Officer Facility & Quality
| 731984 | Donni Edmora | HSC | Panelman II Distill. & Separation
| 747303 | Dwi Fajar Ritonga | HSC | Operator II Reforming
| 10017502 | Eddy Soedjoko | HSC | Sr Operator II Distill. & Separation
| 10015214 | Edward Sitohang | HSC | Sr Operator II Reforming
| 751106 | Erba Syam | HSC | Jr Operator I Reforming
| 750299 | Eri Fradikta | HSC | Operator I Reforming
| 743745 | Erwandie | HSC | Operator II Reforming
| 746462 | Faisal | HSC | Operator II Distill. & Separation
| 746258 | Feirdy Sampurno | HSC | Jr Operator II Reforming
| 750303 | Firman Ikhsan | HSC | Operator I Distill. & Separation
| 695331 | Herbert Maruli Tua | HSC | Panelman II Reforming
| 743778 | Heryanto | HSC | Operator II Reforming
| 748341 | Ibnu Suwondo | HSC | Operator II Reforming
| 714277 | Ibnu Zaenal Arifin | HSC | Section Head HydroSkiming Complex
| 679334 | Ipan Syahrizal | HSC | Sr Supervisor Distill. & Separation
| 10016640 | Johnny Freddy Kaseger | HSC | Officer Facility & Quality
| 10017683 | Machruf | HSC | Shift Supervisor Reforming
| 752817 | Margito | HSC | Jr Operator I Reforming
| 746458 | Mohamad Ridwan | HSC | Operator II Distill. & Separation
| 636882 | Muhamad Sofian | HSC | Panelman II Reforming
| 746247 | Muhammad Agustian Salim | HSC | Jr Operator II Distill. & Separation
| 753515 | Muhammad Zulfikar Lukinanda | HSC | Operator I Distill. & Separation
| 743826 | Oskandar | HSC | Operator II Distill. & Separation
| 743753 | Pajrinur | HSC | Operator II Reforming
| 731951 | Panut Agus Suprihono | HSC | Sr Supervisor Process Reforming
| 10018050 | Paulus Pehang | HSC | Sr Operator II Distill. & Separation
| 751477 | Rifqi Iqbaal Saputra | HSC | Jr Operator I Reforming
| 695218 | Riyanto Pranoto | HSC | Sr Operator II Reforming
| 751479 | Rizki Febrian Akbar | HSC | Jr Operator I Distill. & Separation
| 752749 | Rizky Darmawan | HSC | Jr Operator I Distill. & Separation
| 752838 | Robby Zumar Nuzula Firdaus | HSC | Jr Operator I Reforming
| 752891 | Sabilul Rozak | HSC | Jr Operator I Reforming
| 752834 | Saidin Akhmad Kholid | HSC | Jr Operator I Distill. & Separation
| 636525 | Sandriansyah | HSC | Panelman II Reforming
| 10015000 | Slamet Hariyanto | HSC | Shift Supervisor Reforming
| 743818 | Solikhan | HSC | Operator II Reforming
| 746453 | Sudarno | HSC | Operator II Reforming
| 695201 | Suliyono | HSC | Panelman II Reforming
| 10017011 | Supanto | HSC | Panelman II Distill. & Separation
| 625152 | Supriyanto | HSC | Shift Supervisor Distill. & Separation
| 637238 | Surono | HSC | Shift Supervisor Distill. & Separation
| 637173 | Suroso | HSC | Shift Supervisor Distill. & Separation
| 752808 | Syahruddin | HSC | Jr Operator I Reforming
| 748338 | Syarif Abdul Rahman | HSC | Operator II Distill. & Separation
| 731408 | Togar Parningotan Sinaga | HSC | Panelman II Reforming
| 743801 | Tomi Muhtedi | HSC | Operator II Distill. & Separation
| 637205 | Tronjol Ehsan | HSC | Shift Supervisor Distill. & Separation
| 750374 | Yasin Mulkan Salim | HSC | Operator I Reforming
| 695267 | Yudi Mulyadi | HSC | Sr Operator II Distill. & Separation
| 743842 | Yuliansyah D. | HSC | Operator II Distill. & Separation
| 752047 | Dianuari Kusumawardani | HSE | Jr Engineer I HSE
| 750561 | Martanti Dwi Suryaningtias | HSE | Jr Engineer I HSE
| 718084 | Tjahjo Widojoko | HSE | Manager HSE
| 734187 | Ermawan Agus | Inventory Control | Jr Officer TA & OH Material
| 739177 | Ichsan Devi | Inventory Control | Section Head Inventory Control
| 731757 | M.Abiel F.Y. Kaurong | Inventory Control | Jr Officer MRO RE & NRE
| 731538 | Muhammad Bunyamin | Inventory Control | Supervisor System & Procedure
| 750001 | Poundra Enang Akbar Parusi | Inventory Control | Jr Assistant Procedure & Reporting
| 734073 | Syarwani | Inventory Control | Jr Officer TA & OH Material
| 645524 | Zulkurnain | Inventory Control | Sr Supervisor Inventory Mat. Ops. & Gen.
| 750651 | Fahmi Nur Fauzi | IR & OS Management | Assistant OS Management
| 642795 | I.Ketut Wardana | IR & OS Management | Section Head IR & OS Management
| 751682 | Nurul Salira Utami | IR & OS Management | Junior Analyst Industrial Relation
| 650327 | Suhaidah | IR & OS Management | Senior Supervisor OS Management
| 632937 | Evita Rusdianti | IT RU V | Ast Planning, Budgeting, & Charging
| 731076 | Nasimullah MS | IT RU V | IT RU V - Balikpapan Unit Manager
| 679245 | Abdul Muchid | Laboratory | Sr Tester II Tank & Distribution
| 752835 | Achmad Hari Fabiyanto | Laboratory | Jr Tester I Tank & Distribution
| 625177 | Achmad Maulana | Laboratory | Shift Supervisor CONAL Gas
| 752669 | Achmad Maulana Jabbar | Laboratory | Jr Tester I Tank & Distribution
| 731465 | Agung Kurniawan | Laboratory | Jr Officer Procurement
| 737402 | Akhmad Ikhsan | Laboratory | Sr Tester II Gas
| 731513 | Alex Susanto | Laboratory | Sr Tester I Stream
| 746272 | Ananda Pam Kusdita | Laboratory | Jr Tester II Stream
| 751444 | Andi Ismawan Restu Wirampugi | Laboratory | Jr Tester I Stream
| 679278 | Andrizah Ridhani | Laboratory | Sr Supervisor Feed & Product
| 752675 | Ardin | Laboratory | Jr Tester I NBM
| 750272 | Arizzal Fathul Huda | Laboratory | Tester I Analystic & LL
| 751116 | Arnoldus Ata Koban | Laboratory | Jr Tester I Lab Lawe-lawe
| 642235 | Bambang Irawan | Laboratory | Shift Supervisor Stream & Distribution
| 10018040 | Bambang Suryanto | Laboratory | Supervisor Quality
| 752678 | Bondan Kekey Supriyanto | Laboratory | Jr Tester I NBM
| 679261 | Darmansyah | Laboratory | Sr Tester II NBM
| 748333 | Eddy Supriyadi | Laboratory | Tester II CFR & Aviation
| 731619 | Elvi Sukandar | Laboratory | Sr Tester II CFR & Aviation
| 746310 | Enos Batu Kamban | Laboratory | Jr Tester II Stream
| 748331 | Fahrizal Abdi Rakhman | Laboratory | Tester II Gas
| 687564 | Fatriansyah | Laboratory | Sr Tester II Tank & Distribution
| 752846 | Hardiansyah | Laboratory | Jr Tester I Gas
| 687459 | Hariyanto | Laboratory | Section Head Laboratory
| 679253 | Hendri | Laboratory | Shift Supervisor Stream & Distribution
| 752905 | Herlambang Putranto | Laboratory | Jr Tester I Crude Oil
| 10016821 | Hery Purnadi | Laboratory | Sr Tester II CFR & Aviation
| 639482 | Ilhamto | Laboratory | Sr Tester II Gas
| 679197 | Jamiat | Laboratory | Sr Tester II Crude Oil
| 679115 | Juanda | Laboratory | Sr Tester II Analystic & LL
| 679148 | Kamarudin | Laboratory | Sr Tester II NBM
| 668756 | Kristiawan | Laboratory | Supervisor Facility & Material
| 10018809 | Mohamad Saleh | Laboratory | Shift Supervisor CONAL Gas
| 731635 | Muchamad Zein Azis | Laboratory | Shift Supervisor CONAL Gas
| 748342 | Muh Yusuf Natsir | Laboratory | Tester II Lab Lawe-lawe
| 746254 | Muhammad Danu Rifani | Laboratory | Jr Tester II Stream
| 751471 | Muhammad Eggy Bestari | Laboratory | Jr Tester I Stream
| 750333 | Muhammad Saiful Al Has | Laboratory | Tester I Analystic & LL
| 745252 | Noor Fitriansyah | Laboratory | Tester I CFR & Aviation
| 752676 | Nur Ahmad Abdi | Laboratory | Jr Tester I Crude Oil
| 746255 | Nur Ahmad Zarkasi | Laboratory | Jr Tester II Gas
| 748330 | Oxy Hesmu Prasojo | Laboratory | Tester II Tank & Distribution
| 746303 | Puji Eko Prasetyo | Laboratory | Jr Tester II Stream
| 752671 | Rahmat Kamilin Jufri | Laboratory | Jr Tester I Lab Lawe-lawe
| 751478 | Rinko Angga Wiryanto | Laboratory | Jr Tester I Tank & Distribution
| 753392 | Rizadz Ahmad Syahlie | Laboratory | Tester I Crude Oil
| 751483 | Septi Arie Nur Azizi | Laboratory | Jr Tester I Analystic & LL
| 668626 | Soetoyo | Laboratory | Shift Supervisor CONAL Gas
| 674993 | Suhaimi | Laboratory | Jr Analyst IA & Man Rev
| 10015743 | Sulistya Nugraha | Laboratory | Sr Tester II NBM
| 747302 | Sumarno | Laboratory | Tester II Gas
| 752844 | Sunarko | Laboratory | Jr Technician I Instrum. Analy & Calib.
| 731602 | Suryo Cahyono | Laboratory | Sr Tester II Stream
| 752818 | Sutarjo Rivai | Laboratory | Jr Tester I Lab Lawe-lawe
| 10018379 | Suwaryo | Laboratory | Shift Supervisor Stream & Distribution
| 687491 | Suyanto | Laboratory | Sr Tester II Tank & Distribution
| 642032 | Suyitno | Laboratory | Shift Supervisor Stream & Distribution
| 679131 | Yusran M. | Laboratory | Sr Tester II Analystic & LL
| 10017513 | Yusup M. | Laboratory | Sr Tester II Crude Oil
| 747444 | Adityo Haryoseto | Legal Counsel | Legal Counsel Dispute Management
| 744581 | Anton Sumartono Raharjo | Legal Counsel | Area Manager Legal Counsel RU V
| 753270 | Potayanda | Legal Counsel | Jr Legal Counsel Dispute Management
| 679229 | Tamaji | Legal Counsel | Legal Counsel Contract Mgt
| 747417 | Abnu Musthofa | MA 1 | Jr Technician II Instrument
| 748368 | Affandi Santoso | MA 1 | Technician II Electrical
| 10018375 | Agoes Soewarto | MA 1 | Sr Supervisor Electrical
| 749962 | Ashari Wisnu Wardhana | MA 1 | Technician I Rotating Equip
| 750283 | Bintang Paska Jumrianto Hutasoit | MA 1 | Technician I Instrument
| 752825 | Elman Dahriansyah Nor | MA 1 | Jr Technician I Rotating Equip
| 748367 | Joyo Pratikno | MA 1 | Technician II Mechanical
| 752796 | Linggom Situmorang | MA 1 | Jr Technician I Mechanical
| 752816 | May Helfiansyah | MA 1 | Jr Technician I Mechanical
| 752900 | Mochamad Irawan | MA 1 | Jr Technician I Electrical
| 752830 | Muhammad Rifky Sofiannudin | MA 1 | Jr Technician I Electrical
| 752813 | Mustafa | MA 1 | Jr Technician I Instrument
| 748352 | Norman Sucipto | MA 1 | Technician II Electrical
| 719583 | Paulus Soerjanto | MA 1 | Section Head Maintenance Area 1
| 642227 | Purnomo | MA 1 | Sr Technician II Rotating Equip
| 10016296 | Reins Kobis | MA 1 | Sr Supervisor Instrument
| 694668 | Rumbungan Tambunan | MA 1 | Sr Technician II Mechanical
| 10015003 | Setyo Belo Haryanto | MA 1 | Sr Technician II Mechanical
| 752755 | Wahyu Satrio Wicaksono | MA 1 | Jr Technician I Rotating Equip
| 10018628 | Wayan Sunaka | MA 1 | Sr Supervisor Stationary
| 694716 | Zainal Hakim | MA 1 | Sr Technician II Electrical
| 752897 | Andi Pratama Putra | MA 2 | Jr Technician I Electrical
| 10016105 | Andi Rani | MA 2 | Sr Technician II Mechanical
| 752752 | Beni Santoso | MA 2 | Jr Technician I Electrical
| 752892 | Dicky Muhammad N. | MA 2 | Jr Technician I Rotating Equip
| 10017493 | Elmi | MA 2 | Sr Supervisor Instrument
| 694521 | Fauzi | MA 2 | Sr Supervisor Rotating Equipment
| 752819 | Harsito | MA 2 | Jr Technician I Mechanical
| 748365 | Hasruddin Annas | MA 2 | Technician II Instrument
| 747419 | Ibnu Rahmad | MA 2 | Jr Technician II Instrument
| 747418 | Iryan Budiarsa | MA 2 | Jr Technician II Rotating Equip
| 10016816 | Mohammad Djupri | MA 2 | Sr Supervisor Electrical
| 748359 | Muhamad Kaspuddin | MA 2 | Technician II Electrical
| 752745 | Muhammad Asmi Daryanto | MA 2 | Jr Technician I Instrument
| 752815 | Muhammad Bahrudin | MA 2 | Jr Technician I Rotating Equip
| 748363 | Muhammad Syarif | MA 2 | Technician II Instrument
| 10018816 | Muslimin | MA 2 | Sr Technician II Rotating Equip
| 737638 | Singgih Rahmanto | MA 2 | Sr Technician II Rotating Equip
| 10018377 | Srinopo | MA 2 | Sr Supervisor Stationary
| 694781 | Supriyono | MA 2 | Sr Technician II Rotating Equip
| 714406 | Tabrani | MA 2 | Section Head Maintenance Area 2
| 750723 | Tedi Gampu | MA 2 | Technician I Mechanical
| 645224 | Zairin Zain | MA 2 | Sr Technician II Electrical
| 10018382 | Abdul Aziz | MA 3 | Sr Technician II Rotating Equip
| 752801 | Andi Idrus Syahrial | MA 3 | Jr Technician I Mechanical
| 748357 | Anggika Wiguna | MA 3 | Technician II Rotating Equip
| 752836 | Ardiansyah | MA 3 | Jr Technician I Rotating Equip
| 752748 | Azwirandy | MA 3 | Jr Technician I Instrument
| 748361 | Bambang Satria Surya | MA 3 | Technician II Rotating Equip
| 748356 | Bayu Erfastianto | MA 3 | Technician II Instrument
| 747415 | Bayu Prabowo | MA 3 | Jr Technician II CIVIL
| 674506 | Dedy Kusnaedi | MA 3 | Sr Supervisor Stationary
| 752893 | Gilang Islami | MA 3 | Jr Technician I Electrical
| 748358 | Heldi Pranata | MA 3 | Technician II Electrical
| 741225 | Iwan Kurniawan | MA 3 | Section Head Maintenance Area 3
| 726621 | Kuswanto | MA 3 | Sr Technician II Rotating Equip
| 752890 | M. Fadzkur Rozi | MA 3 | Jr Technician I Mechanical
| 737598 | Muhammad Suryo | MA 3 | Sr Technician II Mechanical
| 653381 | Nahari | MA 3 | Sr Supervisor Instrument
| 10015748 | Prayitno | MA 3 | Sr Supervisor Electrical
| 752798 | Ridolph Thenu | MA 3 | Jr Technician I Electrical
| 747413 | Rizky Eka Saputra | MA 3 | Jr Technician II Instrument
| 747414 | Rudi Eko Styono | MA 3 | Jr Technician II Instrument
| 752799 | Sunarto | MA 3 | Jr Technician I Electrical
| 10017009 | Suprih | MA 3 | Sr Supervisor Rotating Equipment
| 695145 | Wayan Kadar | MA 3 | Sr Technician II Mechanical
| 750727 | Yonzhet Gallaran | MA 3 | Technician I Mechanical
| 10017007 | Abdullah | MA 4 | Sr Technician II Mechanical
| 747412 | Achmad Dawam | MA 4 | Jr Technician II Rotating Equip
| 747420 | Adi Setyawan Wicaksono | MA 4 | Jr Technician II Electrical
| 752797 | Agus Sarwono | MA 4 | Jr Technician I Rotating Equip
| 10016303 | Almaturidi | MA 4 | Sr Technician II Mechanical
| 749959 | Aqwa Aditiyo | MA 4 | Technician I Mechanical
| 750269 | Aris Pajarwako | MA 4 | Technician I Rotating Equip
| 749961 | Arya Indrayana | MA 4 | Technician I Instrument
| 748360 | Ditya Chandra Putra | MA 4 | Technician II Instrument
| 750700 | Dwi Dany Hariadi | MA 4 | Technician I Mechanical
| 752828 | Febriyanto | MA 4 | Jr Technician I Civil
| 731862 | Helmi | MA 4 | Sr Technician II Shift Maintenance
| 748353 | Hernandy Sigit Nurcahyo | MA 4 | Technician II Electrical
| 694846 | Heru Sugeng Purwito | MA 4 | Sr Technician II Rotating Equip
| 10018383 | Jamel Tambunan | MA 4 | Sr Supervisor Rotating Equipment
| 645346 | Jiwo Subono | MA 4 | Sr Technician II Shift Maintenance
| 747416 | K. Andri Purbo Rumekso | MA 4 | Jr Technician II Electrical
| 642405 | Laham | MA 4 | Sr Supervisor Electrical
| 737573 | Mohammad Said | MA 4 | Sr Technician I Electrical
| 752747 | Muhammad Haqqi Tafrizi | MA 4 | Jr Technician I Electrical
| 10018806 | Munawar Halil | MA 4 | Sr Supervisor Stationary
| 747421 | Royan Firdaus | MA 4 | Jr Technician II Civil
| 752821 | Rusmansyah | MA 4 | Jr Technician I Shift Maintenance
| 720554 | Slamet Triyono | MA 4 | Section Head Maintenance Area 4
| 748366 | Sudarmanto | MA 4 | Technician II Rotating Equip
| 752812 | Sugianto | MA 4 | Jr Technician I Shift Maintenance
| 642698 | Sukristo | MA 4 | Sr Technician II Shift Maintenance
| 625558 | Suparmin | MA 4 | Sr Supervisor Lawe - Lawe Maintenance
| 10018042 | Wahyono | MA 4 | Sr Supervisor Instrument
| 645395 | Widodo S. | MA 4 | Sr Technician II Shift Maintenance
| 625696 | Aji Syaid Farouk El Hakim | Maint. & Execution | Officer HSE Compliance
| 717809 | Masrudy | Maint. & Execution | Manager Maintenance Execution
| 674871 | Yaya Sopyani | Maint. & Execution | Officer HSE Compliance
| 751860 | A Arif Madjid | Marine Region VI | Marine Terminal Safety Inspector
| 744473 | Albertus Anto B S | Marine Region VI | Asistant Manager Marine Operation
| 751145 | Armin Sirua | Marine Region VI | Marine Terminal Safety Inspector
| 625744 | Bachriansyah | Marine Region VI | Officer PQC & Bunker
| 625088 | Bastiansyah | Marine Region VI | Jr. Officer Port Operation
| 608183 | Budi Riyadi | Marine Region VI | Jr. Officer Port Operation
| 632953 | Hajir Sulaeman Ago | Marine Region VI | Asman Plan Control & Fasility
| 751180 | Hariyadi | Marine Region VI | Marine Terminal Safety Inspector
| 641433 | Harold Simanjuntak | Marine Region VI | Marine Region VI Manager
| 644641 | Hartatik | Marine Region VI | Sr Spvr Marine Administration
| 644496 | Lisnawati | Marine Region VI | Officer Shipping Agency & Tanker program
| 639611 | Lordy LH Posumah | Marine Region VI | Officer PQC & Bunker
| 608159 | M Kormen | Marine Region VI | Officer PQC & Bunker
| 747713 | M. Yusuf Arifin | Marine Region VI | Supervisor Marine Operation
| 644633 | Marchie Rachel Korag | Marine Region VI | Officer Shipping Agency & Tanker program
| 751185 | Mardiansyah | Marine Region VI | Marine Terminal Safety Inspector
| 751580 | Mestadi Putra | Marine Region VI | Marine Terminal Safety Inspector
| 644528 | Muhammad Noor | Marine Region VI | Officer PQC & Bunker
| 749737 | Muhammad Suwardi Surya Abdi | Marine Region VI | Jr. Officer Port Operation
| 692342 | Priady Bunoto | Marine Region VI | Dockyard Section Head
| 644471 | Selamet Minik | Marine Region VI | Jr. Officer Port Operation
| 644536 | Sugianto | Marine Region VI | Jr. Officer Port Operation
| 751879 | Anharudin | Medical RU V | Officer Industrial & Medical Care RU V
| 682063 | Asmuni | Medical RU V | Officer Industrial & Medical Care RU V
| 718068 | Agus Waluyo | MPS | Manager Maintenance Planning & Support
| 633269 | Kardiman Raharjo | Occupational Health | Section Head Occupational Health
| 701276 | Nurcahyo Prabowo | Occupational Health | Officer Occupational Health Control
| 746277 | Ade Hamid Saputra | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 746301 | Adi Prasetya Nugroho | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 747796 | Adistia Pradika | Oil Movement | Operator II Pump Housing White
| 753512 | Aditya Nugraha | Oil Movement | Operator I Jetty
| 694862 | Agnia A.S | Oil Movement | Supervisor Facility & Quality
| 750255 | Ahmad Najemuddin | Oil Movement | Operator I Pump Hs,1,8,O,P,7,2
| 747306 | Akhmad Fadillah | Oil Movement | Operator II Pump Hs,1,8,O,P,7,2
| 694927 | Alexander R.Tampubolon | Oil Movement | Loading Master Balikpapan Terminal
| 10016819 | Ali Imron | Oil Movement | Sr Supervisor Balikpapan Terminal
| 746308 | Amrin Abdullah | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 746268 | Andhika Eka Putra | Oil Movement | Jr Operator I Jetty
| 10016102 | Anwar | Oil Movement | Shift Supervisor Jetty Terminal
| 10017495 | Ardiansyah | Oil Movement | Shift Supervisor Jetty Terminal
| 746263 | Bagus Yudharta | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 687531 | Bakir | Oil Movement | Sr Operator II Tank
| 694546 | Bambang Purwianto | Oil Movement | Shift Supervisor Lawe - Lawe Terminal
| 746262 | Bambang Riyanto | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 752824 | Bayu Irawan | Oil Movement | Jr Operator I Pump Housing White
| 645598 | Charles Arief Nahar | Oil Movement | Loading Master Balikpapan Terminal
| 10016301 | Darmawan Indra | Oil Movement | Sr Supervisor Lawe - Lawe Terminal
| 746264 | Darul Muslim | Oil Movement | Jr Operator II Jetty
| 751111 | Darwis Mulasit | Oil Movement | Jr Operator I Pump Housing Black
| 714803 | Dede Haerudin | Oil Movement | Section Head Oil Movement
| 746286 | Deni Ilham | Oil Movement | Jr Operator II UTL
| 689013 | Djoko Sulistiyanto | Oil Movement | Sr Operator II Pump Housing White
| 626157 | Djumansyah | Oil Movement | Shift Supervisor Lawe - Lawe Terminal
| 625988 | Djunaidi Kusnadi | Oil Movement | Sr Operator II Pump Housing White
| 636777 | Edy Kartono | Oil Movement | Panelman II Metering System & ATG
| 10017006 | Edyono | Oil Movement | Sr Operator II Pump Housing White
| 668812 | Ery Hertriono | Oil Movement | Loading Master Balikpapan Terminal
| 751458 | Fajar Pratama Himawan | Oil Movement | Jr Operator I Jetty
| 746280 | Hariadi Narhad | Oil Movement | Jr Operator II Jetty
| 10015742 | Hasmi | Oil Movement | Shift Supervisor North Tank Farm
| 750308 | Hendra Saputra | Oil Movement | Operator I Pump Hs,1,8,O,P,7,2
| 752743 | Hendyk Siswanto | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 746304 | Herryanto | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 748332 | Hery Triyono | Oil Movement | Panelman I Metering System & ATG
| 10016642 | Irawan | Oil Movement | Officer Farm & Storage
| 725471 | Iwan Iskandar | Oil Movement | Sr Operator II Pump Hs,5,R & Metering
| 688999 | Jaimansyah | Oil Movement | Loading Master Balikpapan Terminal
| 642098 | Jodis Nainggolan | Oil Movement | Loading Master Balikpapan Terminal
| 751466 | Johan Berliando Sinaga | Oil Movement | Jr Operator I Jetty
| 731554 | Joko Hariyanto | Oil Movement | Loading Master Balikpapan Terminal
| 645662 | Jumikan | Oil Movement | Shift Supervisor South Tank Farm
| 746279 | Jumingun | Oil Movement | Jr Operator II Jetty
| 746271 | Khalel | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 10017003 | Kisman | Oil Movement | Shift Supervisor South Tank Farm
| 746289 | Kurniawan Wahyu Bachtiar | Oil Movement | Jr Operator II Jetty
| 10018060 | Kusaeni | Oil Movement | Shift Supervisor North Tank Farm
| 689038 | Lilik Eko Mulyono | Oil Movement | Officer Jetty
| 737621 | Margono | Oil Movement | Loading Master Balikpapan Terminal
| 695307 | Mariyadi | Oil Movement | Panelman II Metering System & ATG
| 746270 | Marwan Fachrizal | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 746299 | Mohamad Soegibasriko | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 751469 | Mohammad Farid Fathanu | Oil Movement | Jr Operator I Jetty
| 751472 | Muhammad Fadlan Wildan Pranoto | Oil Movement | Jr Operator I Jetty
| 743875 | Muhammad Haryadi | Oil Movement | Operator II Pump Hs,5,R & Metering
| 752245 | Muhammad Muin | Oil Movement | Jr Operator I Pump HS, 5, R & Metering
| 752837 | Muhammad Rajib Mahatir | Oil Movement | Jr Operator I Tank
| 753514 | Muhammad Ramadhan | Oil Movement | Operator I Pump Hs,5,R & Metering
| 641206 | Mujiono | Oil Movement | Panelman II Metering System & ATG
| 732048 | Mumiddin | Oil Movement | Loading Master Lawe - Lawe Terminal
| 689005 | Nahrul Fajri | Oil Movement | Panelman II Metering System & ATG
| 752672 | Naufal Fadhil | Oil Movement | Jr Operator I Jetty
| 753398 | Nugroho Satrio Utomo | Oil Movement | Operator I Tank
| 751473 | Oki Bagus Prasetiyo | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 645581 | Pariyono Setyo A. | Oil Movement | Shift Supervisor Jetty Terminal
| 752810 | Purwanto | Oil Movement | Jr Operator I Tank
| 746265 | Raden Muhammad Yanuar Zain | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 746295 | Rahmadi | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 751474 | Rebo | Oil Movement | Jr Operator I UTL
| 751475 | Reza Fitrian | Oil Movement | Jr Operator I Pump Housing White
| 751476 | Rian Kurniawan | Oil Movement | Jr Operator I Pump Housing Black
| 752744 | Ricky Azmi Zain | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 746293 | Riswandi Syam | Oil Movement | Jr Operator II UTL
| 751480 | Rohmad Bagus Hermawan | Oil Movement | Jr Operator I Pump Housing Black
| 745259 | Rudi Ariston Maruba Hutapea | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 751481 | Rudiansyah | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 746273 | Rudianto | Oil Movement | Jr Operator II Pump Hs,5,R & Metering
| 624845 | Rusdani | Oil Movement | Loading Master Balikpapan Terminal
| 689095 | Said Hanafiah | Oil Movement | Panelman II Metering System & ATG
| 10017679 | Salmansyah | Oil Movement | Loading Master Balikpapan Terminal
| 750354 | Sandi Sarira | Oil Movement | Operator I Pump Hs,1,8,O,P,7,2
| 746260 | Saparudin | Oil Movement | Jr Operator II Pump Hs,1,8,O,P,7,2
| 10016302 | Sarminto | Oil Movement | Shift Supervisor North Tank Farm
| 10018049 | Siswoyo | Oil Movement | Shift Supervisor North Tank Farm
| 743834 | Solichan | Oil Movement | Operator II UTL
| 645232 | Sony Fadilah | Oil Movement | Shift Supervisor Lawe - Lawe Terminal
| 746285 | Subari Tambung | Oil Movement | Jr Operator II UTL
| 637132 | Subarkah | Oil Movement | Sr Operator II Pump Hs,5,R & Metering
| 752826 | Suharno | Oil Movement | Jr Operator I Pump Hs,1,8,O,P,7,2
| 645573 | Sujarwo | Oil Movement | Loading Master Balikpapan Terminal
| 752827 | Sujoko | Oil Movement | Jr Operator I Pump Hs,1,8,O,P,7,2
| 10016643 | Sultan | Oil Movement | Sr Operator II Pump Housing White
| 10014998 | Sunaryo | Oil Movement | Sr Operator II Pump Housing White
| 645646 | Suryatno | Oil Movement | Loading Master Balikpapan Terminal
| 645679 | Sutoyo | Oil Movement | Sr Operator II Pump Housing White
| 752243 | Sutriono | Oil Movement | Jr Operator I Pump Hs,5,R & Metering
| 746249 | Sutrisno | Oil Movement | Jr Operator II Pump Housing White
| 645654 | Suwito | Oil Movement | Panelman II Metering System & ATG
| 746297 | Syarifuddin | Oil Movement | Jr Assistant System Database
| 645621 | Tames Pandairot | Oil Movement | Sr Operator II Pump Housing White
| 10014997 | Totok Widarto | Oil Movement | Shift Supervisor South Tank Farm
| 695031 | Tulus Hernowo | Oil Movement | Sr Operator II Pump Hs,5,R & Metering
| 642802 | Untung Sutikno | Oil Movement | Loading Master Lawe - Lawe Terminal
| 10016305 | Wahyono | Oil Movement | Shift Supervisor South Tank Farm
| 751110 | Wahyu Hendryanto | Oil Movement | Jr Operator I Pump Housing White
| 751117 | Yulius | Oil Movement | Jr Operator I Pump Housing Black
| 746315 | Yuri Susanda | Oil Movement | Jr Operator II Jetty
| 611269 | Zainul Rachman | Oil Movement | Shift Supervisor Jetty Terminal
| 748334 | Zulkifli | Oil Movement | Panelman I Metering System & ATG
| 714788 | Ahmad Syopian | OPI | Sr Analyst Workstream Reliability
| 746562 | Doan Herdani | OPI | Analyst Workstream Optimization
| 611041 | Eddy Suryamadi | OPI | Sr Analyst Port Integration Network
| 708259 | Hermawan Bayu | OPI | Manager OPI
| 720287 | Joko Pitoyo | OPI | Sr Analyst Workstream Refinery HSE
| 636566 | Marsudi Widodo | OPI | Analyst Workstream Refinery HSE
| 624326 | Norhan Effendi | OPI | Analyst Workstream Refinery Loss
| 744703 | Riny Utami | OPI | Sr Analyst Workstrem,MindCap,Lead&MgtSys
| 10018380 | Surya Noor Ananta | OPI | Sr Analyst Workstream Refinery Fuel
| 10017705 | Fitri Nurlia | Opr. & Manufacturing | Staf Utama Diperbantukan
| 717477 | Gunarno | Opr. & Manufacturing | Sr Manager Operation & Manufacturing
| 10016098 | Abdullah | Planning & Schedulling | Sr Supervisor Work Plan & Budget
| 747504 | Adi Nugroho | Planning & Schedulling | Engineer Electrical Planning
| 650279 | Andang Iskandar | Planning & Schedulling | Sr Supervisor Planning
| 747525 | David P Nainggolan | Planning & Schedulling | Engineer Rotating Equip. Planning
| 633228 | Irianto | Planning & Schedulling | Engineer Civil & Building Planning
| 719534 | Ismed | Planning & Schedulling | Section Head Planning & Scheduling
| 747878 | Kurniawan Dahri | Planning & Schedulling | Engineer PM & CM Planning
| 10016300 | M. Agus Setyabudi | Planning & Schedulling | Engineer Rotating Equip. Planning
| 749524 | Mayapati | Planning & Schedulling | Engineer Civil & Building Planning
| 645127 | Munir M. | Planning & Schedulling | Engineer Mechanical Planning
| 749834 | Pancar Fransco | Planning & Schedulling | Jr Engineer II Planning
| 737565 | Parji | Planning & Schedulling | Sr Supervisor Maintenance System
| 747544 | Ronald Iskandar Hanafi | Planning & Schedulling | Engineer Instrument Planning
| 747545 | Ryan Arianto | Planning & Schedulling | Engineer Mechanical Planning
| 10015747 | Sofiantono | Planning & Schedulling | Engineer Civil & Building Planning
| 608694 | Sugeng Sulistyono | Planning & Schedulling | Engineer Instrument Planning
| 674563 | Dwi Wahyuningsih | Plant Reliability | Engineer New Plant / Mod Reliability
| 608256 | Syaifullah | Plant Reliability | Engineer Sec. Proc. Plant Reliability
| 719704 | Teguh Puji Handoko | Plant Reliability | Section Head Plant Reliability
| 752540 | Abdullah Al Mukhsi | Process Eng | Jr Engineer I Primary Process
| 752529 | Agustin Wulandari | Process Eng | Jr Engineer I Secondary Process
| 751800 | Agya Kumala Asri | Process Eng | Jr Engineer I Primary Process
| 750538 | Ahmad Banadib | Process Eng | Jr Engineer I Primary Process
| 752535 | Andre Yanda Putra | Process Eng | Jr Engineer I Primary Process
| 748866 | Andri Hariagung | Process Eng | Jr Engineer II Primary Process
| 750540 | Angelin Ayu Palupi | Process Eng | Jr Engineer I Primary Process
| 751805 | Aries Purijatmiko | Process Eng | Jr Engineer I Secondary Process
| 750920 | Arini Aristia Saputra | Process Eng | Jr Engineer I Secondary Process
| 750547 | Dani Wibowo | Process Eng | Jr Engineer I Primary Process
| 752531 | Dede Setiadi | Process Eng | Jr Engineer I Process Control
| 749815 | Didiet Tri Hardoyo | Process Eng | Jr Engineer II Primary Process
| 741144 | Djatmiko Darmo Saputro | Process Eng | Lead of Process Engineering
| 748639 | Erick Leonardo | Process Eng | Jr Engineer II Secondary Process
| 748871 | Erwan Yonata | Process Eng | Jr Engineer II Primary Process
| 747889 | Ery Gunarto | Process Eng | Engineer CDU/LPGRec/SWS/LPG Treat Proc
| 748640 | Fajri Kautsar | Process Eng | Jr Engineer II Primary Process
| 748873 | Febriar Cahyaratri | Process Eng | Jr Engineer II Secondary Process
| 747485 | Ihya Ramadhanu Khairil | Process Eng | Engineer CDU V/HVU-III/Wax Plant Proc
| 625322 | Ilham | Process Eng | Group Leader Primary Process
| 746926 | Julius Pasaribu | Process Eng | Engineer Process Safety
| 748651 | Marosbi Lamasta | Process Eng | Jr Engineer II Secondary Process
| 750928 | Misbaqul Munir | Process Eng | Jr Engineer I Secondary Process
| 752533 | Muchlis Riki Dharmawan | Process Eng | Jr Engineer I Process Control
| 679383 | Muchsin M. | Process Eng | Engineer Environment
| 747896 | Mu'ti Fahmi | Process Eng | Engineer Boil/TG/Demin Pl/SWD-CW I Proc
| 750585 | Nugraha Thariq Majid | Process Eng | Jr Engineer I Secondary Process
| 748659 | Pungki Prayuda Ferdian | Process Eng | Jr Engineer II Primary Process
| 750930 | Raffael Reyno Raditya | Process Eng | Jr Engineer I Primary Process
| 751832 | Ragaguci | Process Eng | Jr Engineer I Primary Process
| 748887 | Rano Indra Pamungkas YP | Process Eng | Jr Engineer II Primary Process
| 747899 | Rendra Jayantara Putra | Process Eng | Engineer APC - DCS
| 746582 | Roland Joshua Panjaitan | Process Eng | Group Leader Secondary Process
| 748893 | Shokma Perdana Adhigraha | Process Eng | Jr Engineer II Secondary Process
| 750571 | Suharsono | Process Eng | Jr Engineer I Secondary Process
| 748669 | Taufik Basuki Hartopo | Process Eng | Jr Engineer II Secondary Process
| 748670 | Tofan Santoso | Process Eng | Jr Engineer II Secondary Process
| 746588 | Widakdo | Process Eng | Group Leader Process Control
| 749844 | Yohanes Indra Wibisana | Process Eng | Jr Engineer II Secondary Process
| 751844 | Yudhistiro Doddy Sadewo | Process Eng | Jr Engineer I Secondary Process
| 706096 | Chatarina Duma | Procurement | Manager Procurement
| 624342 | Ali Mansur | Production | Shift Superintendent
| 624878 | Bambang Margono | Production | Shift Superintendent
| 731335 | Harry Siswaji | Production | Officer HSE Comp
| 679918 | Naipospos H. | Production | Officer HSE Comp
| 679715 | Saiful Alamsyah | Production | Officer HSE Comp
| 713645 | Sugeng Firmanto | Production | Manager Production
| 624415 | Suwito Widodo | Production | Shift Superintendent
| 679837 | Darmawansyah | Project Eng | Jr Engineer II Cost Control & Estimate
| 746929 | Darto Martua Silitonga | Project Eng | Engineer Instrument & Control Project
| 694692 | Dullen Arkadius | Project Eng | Engineer Electrical Project
| 714682 | Handi Soejanto Ateng | Project Eng | Section Head Project Engineering
| 746579 | Muttaqin Showwabi | Project Eng | Sr Supervisor Cost. Eng. & Project Ctrl.
| 737557 | Nuryanto | Project Eng | Jr Engineer II Mechanical, RE & Civil
| 737581 | Saptono Nugroho | Project Eng | Jr Engineer II Electrical & Instrument
| 679529 | Sulardi | Project Eng | Engineer Civil Project
| 10015217 | Suwondo | Project Eng | Sr Supervisor Construction Project
| 650798 | Aisah Nurmayanti | Purchasing | Sr Supervisor Purchasing Mat. Ops.& Gen.
| 747509 | Fifi Andriyani Agustin | Purchasing | Buyer RE Material
| 695315 | Hadi Setya Iswara | Purchasing | Buyer TA Material
| 694798 | Herman Nibras | Purchasing | Buyer General Material
| 726557 | Kuwat Imam Budiono | Purchasing | Buyer Program Material
| 720521 | Oktoriza | Purchasing | Section Head Purchasing
| 734138 | Rakhman | Purchasing | Buyer Electrical & Instrument
| 750010 | Rifan Nasrul Arifin | Purchasing | Jr Assistant Administration Purchasing
| 653462 | Siswoyo | Purchasing | Buyer TA Material
| 10017005 | Sudaryanto | Purchasing | Sr Supervisor Purchasing Material Maint.
| 694838 | Farid Rawung | Quality Management | Section Head Quality Management RU V
| 737387 | Hurianto | Quality Management | Analyst Quality Management
| 733093 | Mohammad Zulfihar | Quality Management | Analyst Quality Management
| 756721 | Ferry Hadianto | Quantity Accounting Kantor Pusat
| 747867 | Tasha Armanda Tene | Quantity Accounting Kantor Pusat
| 739988 | Rusdi Rachmani | Ref. Planning & Opt. | Manager Refinery Planning & Optimization
| 746552 | Alex Wahyu Rahadi | Refinery Planning | Sr Supervisor Plann. Opt. Intake & Prod.
| 720319 | M. Alaiddin | Refinery Planning | Section Head Refinery Planning
| 747309 | Rusli | Refinery Planning | Jr Assistant Data Linear Programming
| 748674 | Wulan Novitasari | Refinery Planning | Jr Analyst Select Crude & Intermedia
| 718043 | Agus Maulana | Reliability | Manager Reliability
| 720481 | Lam Asi Lumban Toruan | Reliability | Sr Engineer Instr. Equipment Reliability
| 751715 | Adi Sulistiyono | RIA | Jr Auditor Internal Audit Kalimantan
| 749337 | Edziardo Edazmara | RIA | Auditor Internal Audit Kalimantan
| 748825 | Harry Valentine Turnip | RIA | Auditor Internal Audit Kalimantan
| 752555 | Nabilla Tamimi | RIA | Jr Auditor Internal Audit Kalimantan
| 744683 | Nahdiyani Hasbi | RIA | Internal Audit Kalimantan Manager
| 751727 | Septika Soputri Silaban | RIA | Jr Auditor Internal Audit Kalimantan
| 750799 | Unang Wibowo | RIA | Sr Auditor Internal Audit Kalimantan
| 751729 | Zalina Hayati | RIA | Jr Auditor Internal Audit Kalimantan
| 747870 | Achmad Dwi Nugroho | Rot. Equip. Insp. Engineer | Engineer Rotating Con Mon Inspection
| 751803 | Arandityo Narutomo | Rot. Equip. Insp. Engineer | Jr Engineer I Rotating Equip Inspection
| 714171 | Beben Bunyamin | Rot. Equip. Insp. Engineer | Lead of RE Inspection Engineer
| 10015744 | Budiono | Rot. Equip. Insp. Engineer | Group Leader Rotating Equip Inspection
| 749814 | Chandra Ricardo Marulitua Simanjuntak | Rot. Equip. Insp. Engineer | Jr Engineer II Rotating Equip Inspection
| 747982 | Farid Rakhmad Wijaya | Rot. Equip. Insp. Engineer | Engineer Rotating Equip. Inspection
| 10017698 | Suparno | Rot. Equip. Insp. Engineer | Engineer Rotating Equip. Inspection
| 749845 | Yoppi Sutrisna | Rot. Equip. Insp. Engineer | Jr Engineer II Rotating Equip Inspection
| 747562 | Adi Firmansyah | Safety | Sr Supervisor PSM, Regulation & Standard
| 748642 | Farida Hayati | Safety | Jr Officer PSM & System Regulation
| 731887 | Kukuh Santoso | Safety | Jr Officer Process Safety Dis Wax
| 720538 | Oky Wibisono | Safety | Section Head Safety
| 753470 | Roni Harimuljana Zulkarnain | Safety | Sr Supervisor Safety Advisor
| 725277 | Sugeng Priambodo | Safety | Jr Officer Process Safety UTL
| 750088 | Suroto Budi Prakoso | Safety | Jr Officer Process Safety Offsite & OM
| 714747 | Heru Subekti | Sche. Mat. & Svc. Support | Section Head Sched.Mat.&Service Support
| 645484 | Jhonny Amir | Sche. Mat. & Svc. Support | Analyst Program Schedulling
| 750076 | Arjamal Waginopo | Security | Group Head Housing & Non Refinery Area
| 750094 | Asep Cahyana | Security | Sr Supervisor Physical Of Security
| 753368 | Asto Pranoto | Security | Group Head South Area
| 628888 | Eddy Syahfiarly | Security | Group Head Refinery Lawe-Lawe Area
| 748511 | Edy Eko Iriantoso | Security | Sr Supervisor Non Physical Security
| 753360 | Jemingan | Security | Group Head Refinery Area
| 629073 | Muhammad Anwar | Security | Jr Officer Investigation & Evaluation
| 748512 | Muhammad Nur | Security | Assistant Secure Identification Personal
| 750082 | R. A. Prananto Akoso | Security | Section Head Security
| 750072 | Sucipto | Security | Group Head Refinery & Sei Wein Area
| 705537 | Sugiyanta | Security | Sr Supervisor Secure Legal.IP,Bgt.Ad
| 628911 | Sukadi | Security | Jr Officer Investigation & Evaluation
| 679667 | Suyitno | Security | Jr Officer Intelligence
| 752566 | Aprian Dwi Putra | Services and Warehousing | Assistant General Material Warehouse
| 749969 | Dedy Asman | Services and Warehousing | Jr Assistant Formality & Administration
| 750702 | Fahrurrozi Manalu | Services and Warehousing | Jr Assistant Housing Services
| 749997 | Mesakh Yudi Saputra | Services and Warehousing | Jr Assistant Operational & Pool
| 737646 | Mohammad Ilham | Services and Warehousing | Assistant TA & OH Warehouse
| 608807 | Murdi | Services and Warehousing | Section Head Services & Warehousing
| 656638 | Pomantow M.E.N. | Services and Warehousing | Assistant Chemical & Catalyst Warehouse
| 750721 | Sunanthi Patria Wijaya | Services and Warehousing | Jr Assistant Office Services
| 750724 | Teo Aprianto Palondongan | Services and Warehousing | Jr Assistant Housing Services
| 651989 | Warouw Arthur | Services and Warehousing | Sr Supervisor Receiving
| 752006 | Agung Tri Laksana | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 753243 | Ahmad Nurdin Arpah | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 750579 | Aji Mahardika | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 747522 | Apri Marlinaldi Majid | Stat. & Statu. Insp. Engineer | Engineer Corrosion Metallurgy Inspection
| 751808 | Betcylya Br Sibarani | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 751810 | Budi Utomo | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 645338 | Djuwarto | Stat. & Statu. Insp. Engineer | Engineer Stationary Inspection
| 749973 | Dwi Yuda Agus Mulyadi | Stat. & Statu. Insp. Engineer | Technician I Stationary Inspection
| 749983 | Hendry Sudarma | Stat. & Statu. Insp. Engineer | Technician I Stationary Inspection
| 10016299 | Ibramsyah | Stat. & Statu. Insp. Engineer | Engineer Statutory Inspection
| 747536 | Lis Sugiantoro | Stat. & Statu. Insp. Engineer | Engineer Stationary Inspection
| 10018376 | Muhamad Saleh | Stat. & Statu. Insp. Engineer | Group Leader Statutory Inspection.
| 663693 | Nana Sukarna | Stat. & Statu. Insp. Engineer | Group Leader Stationary Inspection
| 748354 | Nur Irpansyah | Stat. & Statu. Insp. Engineer | Technician II Stationary Inspection
| 679764 | Poseng Silaen | Stat. & Statu. Insp. Engineer | Engineer Tank Farm & Infras. Inspection
| 720165 | Rahendrafedy | Stat. & Statu. Insp. Engineer | Lead of Station.& Statu. Insp. Engineer
| 750795 | Samsul Arifin | Stat. & Statu. Insp. Engineer | Engineer Stationary Inspection
| 750570 | Septyan Dwi Adrianto | Stat. & Statu. Insp. Engineer | Jr Engineer I Stationary Inspection
| 625509 | Suko Rahardjo | Stat. & Statu. Insp. Engineer | Engineer Stationary Inspection
| 748344 | Dicky Prasetya | Supply Chain & Distribution | Jr Assistant Transaction Product
| 10017496 | Marsudi Yana | Supply Chain & Distribution | Sr Supervisor Supply Chain Product
| 720335 | Nur Fauzah | Supply Chain & Distribution | Section Head Supply Chain & Distribution
| 748657 | Prasojo Dwi Cahyopramono | Supply Chain & Distribution | Jr Analyst Supply Chain Crude Import
| 731302 | R.Ucu Allpaeruzzabadi | Supply Chain & Distribution | Jr Analyst Commercial Product
| 10017488 | Sjaiful Islam | Supply Chain & Distribution | Sr Supervisor Crude & Process Intermedia
| 10016099 | A.Hanafi S | T/A | Engineer Electrical T/A Planning
| 720076 | Agung Prijadi | T/A | Section Head Turn Around
| 695178 | Buchairi Zaman | T/A | Engineer I Mechanical T/A Planning
| 10016100 | Sigit Budiardjo | T/A | Engineer Civil T/A Planning
| 694765 | Sugeng Kusmoro | T/A | Engineer II Mechanical T/A Planning
| 642081 | Sugeng Riyanto | T/A | Engineer II Mechanical T/A Planning
| 750766 | Wisnu Purwanto | T/A | Jr Engineer II Instrument T/A Planning
| 708104 | I Gusti Agung Nyoman Bismayudha | Turn Around | Manager Turn Around
| 739906 | Bambang Harimurti | Unit HR RU V | RDMP Project Team
| 708201 | Imam Sunarto | Unit HR RU V | RDMP Project Team
| 713637 | Jan Widi Widodo L. | Unit HR RU V | RDMP Project Team
| 749166 | Ragil C. Pamungkas | Unit HR RU V | RDMP Project Team
| 714585 | Rully Hendarsetiawan | Unit HR RU V | RDMP Project Team
| 719145 | Susapto Yuwono | Unit HR RU V | Unit Manager HR RU V
| 717777 | Syaifuddin Azhar | Unit HR RU V | RDMP Project Team
| 752803 | Aan Padli | Utilities | Jr Operator I DEMIN
| 747097 | Abdul Rachman Saleh | Utilities | Operator II Boiler
| 751436 | Achmad Rifai | Utilities | Jr Operator I WTP
| 747795 | Achmad Rohimin | Utilities | Operator II Shift Distrib.Steam & Air
| 746269 | Aditya Christy Rosso Rindo R. | Utilities | Jr Operator II Turbin Generator
| 746266 | Agung Kurniadi | Utilities | Jr Operator II Turbin Generator
| 751437 | Agung Mas Rupiawan | Utilities | Jr Operator I WTP
| 751438 | Agung Prabasworo | Utilities | Jr Operator I S. WAIN
| 747308 | Agus Siswa Triharjo | Utilities | Operator II Boiler
| 751439 | Agus Supriyadi | Utilities | Jr Operator I Shift Distrib.Steam & Air
| 625963 | Agus Triyanto S. | Utilities | Shift Supervisor DEMIN, SWD & CWI
| 748340 | Agustinus Sonda | Utilities | Operator II Shift Distrib.Steam & Air
| 752894 | Ahmad Ria Saifudin | Utilities | Jr Operator I WTP
| 746283 | Alhaqu Muhamad Murtado | Utilities | Jr Operator II Boiler
| 752677 | Andi Asmier Prawiryo | Utilities | Jr Operator I WTP
| 725803 | Andi Ilyas | Utilities | Shift Supervisor Distribution
| 750260 | Andi Oddang | Utilities | Operator I Shift Distrib.Steam & Air
| 751445 | Andi Zulfadillah | Utilities | Jr Operator I CWI & RPAL
| 750265 | Anri Silalahi | Utilities | Operator I Turbin Generator
| 637116 | Ariansyah | Utilities | Shift Supervisor DEMIN, SWD & CWI
| 746323 | Arif Prabowo | Utilities | Jr Operator II Turbin Generator
| 746322 | Arif Zainuri | Utilities | Jr Operator II CWI & RPAL
| 746325 | Aris Setiyawan | Utilities | Jr Operator II Turbin Generator
| 626084 | Armani | Utilities | Shift Supervisor Boiler
| 746282 | Ashadi | Utilities | Jr Operator II Turbin Generator
| 746290 | Bagus Rochman | Utilities | Jr Operator II Turbin Generator
| 751450 | Bagus Satrianto Hadi | Utilities | Jr Operator I Boiler
| 10016297 | Bimo Wijonarko | Utilities | Shift Supervisor DEMIN, SWD & CWI
| 751454 | Cahyo Agus Maulana | Utilities | Jr Operator I Shift Distrib.Steam & Air
| 746313 | Daniel Perunan Toke | Utilities | Jr Operator II Shift Distrib.Steam & Air
| 731562 | Dedi Fitria Antonova | Utilities | Panelman II Boiler
| 751456 | Dedi Wijaya | Utilities | Jr Operator I Shift Distrib.Steam & Air
| 746294 | Dedy Saputra Gultom | Utilities | Jr Operator II Boiler
| 751457 | Dika Mahendra | Utilities | Jr Operator I Turbin Generator
| 752805 | Donny Surya Putra | Utilities | Jr Operator I WTP
| 746306 | Eliansyah | Utilities | Jr Operator II Turbin Generator
| 746318 | Estu Prayogi | Utilities | Jr Operator II Boiler
| 642705 | Felix Erdy Van Houten | Utilities | Officer Facility & Quality
| 751107 | Gunawan Febri Harlison | Utilities | Jr Operator I SWD
| 645192 | Halimin | Utilities | Panelman II Turbin Generator
| 749980 | Hari Widodo | Utilities | Operator I Boiler
| 746317 | Hartono | Utilities | Jr Operator II Turbin Generator
| 752246 | Herman Bahana | Utilities | Jr Operator I Turbin Generator
| 746573 | I Putu Gede Eka Mahaswara | Utilities | Sr Supervisor Process Water
| 642413 | I.Ketut Warsa | Utilities | Officer Facility & Quality
| 748336 | Impola Roni P. Simatupang | Utilities | Operator II Electrical Operation
| 750316 | Ivan Edo Nurhadist | Utilities | Operator I Turbin Generator
| 745254 | Julianus Walelangi | Utilities | Operator I Shift Distrib.Steam & Air
| 10018815 | Kunto Wibisono | Utilities | Shift Supervisor WTP & S. WAIN
| 642268 | Mahyuddin Suleman | Utilities | Shift Supervisor WTP & S. WAIN
| 752814 | Miswanto | Utilities | Jr Operator I DEMIN
| 750323 | Muh. Nur Hidayat | Utilities | Operator I Boiler
| 746311 | Muhammad Ary Fauzi Kurniawan | Utilities | Jr Operator II Turbin Generator
| 753404 | Muhammad Syarifullah | Utilities | Operator I Turbin Generator
| 10018810 | Mulyanto | Utilities | Panelman II Turbin Generator
| 609114 | Mulyanto.A | Utilities | Panelman II Turbin Generator
| 746309 | Mustafa Mukmin Putra | Utilities | Jr Assistant Material & Data Supporting
| 10018046 | Nur Alamsyah | Utilities | Panelman II Boiler
| 746321 | Nur Hadinata | Utilities | Jr Operator II SWD
| 746287 | Nur Samsudi | Utilities | Jr Operator II Boiler
| 751115 | Nurdin | Utilities | Jr Operator I S. WAIN
| 752899 | Prima Ammar Fauzan | Utilities | Jr Operator I CWI & RPAL
| 746288 | Purnama Putra Pontoh | Utilities | Jr Operator II Boiler
| 10017725 | Rabarno | Utilities | Sr Supervisor Distribution
| 746259 | Raden Iman Subhan Mulyadi | Utilities | Jr Operator II Turbin Generator
| 752742 | Ramadhan Satrio A. | Utilities | Jr Operator I WTP
| 746274 | Rizal Indrawan | Utilities | Jr Operator II Turbin Generator
| 745249 | Rizal Rifai Rusman | Utilities | Operator I Electrical Operation
| 746320 | Salehuddin | Utilities | Jr Operator II Turbin Generator
| 642284 | Salman | Utilities | Shift Supervisor Turbin Generator
| 10016306 | Salyo | Utilities | Shift Supervisor Turbin Generator
| 746251 | Sebnoe Setyo Prayugo | Utilities | Jr Operator II Boiler
| 694724 | Sinolungan A.E | Utilities | Shift Supervisor Distribution
| 636858 | Sudaryatmiko Darmin | Utilities | Panelman II Boiler
| 10018813 | Sudodo | Utilities | Shift Supervisor Distribution
| 746257 | Sugirianto | Utilities | Jr Operator II Boiler
| 10017013 | Suharno | Utilities | Shift Supervisor Boiler
| 746275 | Supangat | Utilities | Jr Operator I Electrical Operation
| 642673 | Supriadi | Utilities | Panelman II Turbin Generator
| 10018613 | Surasono | Utilities | Shift Supervisor Turbin Generator
| 731546 | Suroso | Utilities | Shift Supervisor Boiler
| 746253 | Susanta | Utilities | Jr Operator II Boiler
| 10018614 | Suwari | Utilities | Sr Supervisor Process Generator
| 714666 | Syarif Hidayat | Utilities | Section Head Utilities
| 731879 | Tarmiji | Utilities | Shift Supervisor Distribution
| 752804 | Teguh Prihatin | Utilities | Jr Operator I WTP
| 642316 | Thamrin Rustam | Utilities | Panelman II Turbin Generator
| 752902 | Tito Hartanto | Utilities | Jr Operator I DEMIN
| 10015215 | Unus | Utilities | Shift Supervisor Turbin Generator
| 746319 | Virdan | Utilities | Jr Operator II S. WAIN
| 731481 | Widi Hartanto | Utilities | Shift Supervisor DEMIN, SWD & CWI
| 746316 | Wigih Eko Prabowo | Utilities | Jr Operator II Boiler
| 752832 | Yosandi Ibnu Kholid | Utilities | Jr Operator I SWD
| 747092 | Yusuf | Utilities | Operator II DEMIN
| 746256 | Zainal Abidin | Utilities | Jr Operator II SWD
| 731343 | Zuhri Swidhonarko | Utilities | Officer Facility & Quality
| 745250 | Zulkifli | Utilities | Jr Operator II S. WAIN
| 737613 | Agus Syahroni | Workshop | Sr Technician II Stationary
| 10015570 | Arief Sumedi M. | Workshop | Sr Technician II Machinery & RE
| 737702 | Arijadi Noor | Workshop | Sr Technician II Elect /Inst Lab
| 694813 | Bahtarudin | Workshop | Shift Supervisor Maintenance
| 694732 | Benny Panigoro | Workshop | Sr Technician II Heavy Equip&Tran. Maint
| 694619 | Darwis Siregar | Workshop | Sr Technician II Rigging & Scaffolding
| 10016097 | Donner Nainggolan | Workshop | Sr Supervisor Mechanical Workshop
| 750703 | Fauzan Rahman | Workshop | Technician I Maintenance
| 10017731 | Hasanuddin | Workshop | Shift Supervisor Maintenance
| 656727 | Hotler Napitupulu | Workshop | Sr Technician II Master Tools
| 748362 | Ilham Nahru | Workshop | Technician II Maintenance
| 642243 | Imam Subawai | Workshop | Sr Technician II Electrical
| 695015 | Jaka Sukiyanto | Workshop | Sr Technician II Stationary
| 10018056 | Jonner Hutahaean | Workshop | Sr Technician II Heavy Equip&Tran. Maint
| 694992 | Mantoro | Workshop | Sr Technician II Stationary
| 748355 | Mardian Aqsa | Workshop | Technician II Machinery & RE
| 694902 | Mujiono | Workshop | Shift Supervisor Maintenance
| 608783 | Nur Harianto | Workshop | Sr Supervisor Elect & Inst Workshop
| 10017004 | Paidi Syariah | Workshop | Sr Supervisor Heavy EQP & Rigger
| 719607 | Ricky Paulus J.Langitan | Workshop | Section Head Workshop
| 694943 | Selamat Hutajulu | Workshop | Sr Technician II Rigging & Scaffolding
| 694684 | Suhartono | Workshop | Shift Supervisor Maintenance
| 750369 | Topit Prianto | Workshop | Technician I Maintenance
| 731773 | Yohannis Dwi Purnomo | Workshop | Sr Technician II Master Tools

## Tαnki ROAS (hidden)
TypeTank_FK | Tank_PK | Plant | TypeTank | TankArea | Diameter | Vol_Bersih | Tinggi_Lubang_Ukur | Tinggi_Tanki | Tinggi_Meja | Massa_Jenis | Koefisien_Muai | Faktor_Kor_Vol | Min_Tinggi_Cairan | Max_Tinggi_Cairan | Suhu | Product_FK | ProductName | Tinggi_Drawable | Faktor_Kor_Vol1 | Tanggal_Tera_Ulang | Tanggal_Tera_Akhir | IsActive
2 | 101T1A | R501 | FLOATING ROOF | LAWE2 | 95018 | 128443775 | 21048 | 19503 | 264 | 0.8492 | 3.48e-05 | 36 | 500 | 18800 | 36 | MM | CRUDE OIL,LAWE | 500 | 36 | 9/28/2019 12:00:00 AM | 9/28/2013 12:00:00 AM | Active
2 | 101T1B | R501 | FLOATING ROOF | LAWE2 | 95124 | 128553640 | 21075 | 19504 | 236 | 0.7981 | 3.48e-05 | 76.5 | 0 | 18300 | 38 | MM | CRUDE OIL,LAWE | 500 | 76.5 | 5/31/2019 12:00:00 AM | 5/24/2013 12:00:00 AM | Active
2 | 101T1C | R501 | FLOATING ROOF | LAWE2 | 95043 | 128483898 | 21065 | 19498 | 208 | 0.8438 | 3.48e-05 | 70.2 | 0 | 18300 | 38 | MM | CRUDE OIL,LAWE | 500 | 70.2 | 5/31/2019 12:00:00 AM | 5/22/2013 12:00:00 AM | Active
2 | 101T1D | R501 | FLOATING ROOF | LAWE2 | 95046 | 128474786 | 21083 | 19503 | 225 | 0.8428 | 3.48e-05 | 68.6 | 500 | 18300 | 33 | MM | CRUDE OIL,LAWE | 500 | 68.6 | 2022-01-03 | 2016-01-03 | Active
2 | 101T1E | R501 | FLOATING ROOF | LAWE2 | 95073 | 128580224 | 21064 | 19499 | 251 | 0.7858 | 3.48e-05 | 78.9 | 500 | 18300 | 35 | MM | CRUDE OIL,LAWE | 500 | 78.9 | 9/13/2020 12:00:00 AM | 9/13/2014 12:00:00 AM | Active
2 | 101T1F | R501 | FLOATING ROOF | LAWE2 | 95023 | 127887420 | 20805 | 19511 | 227 | 0.7981 | 3.48e-05 | 30 | 500 | 18340 | 30 | MM | CRUDE OIL,LAWE | 500 | 30 | 2023-11-11 | 2017-11-12 | Active
2 | 101T1G | R501 | FLOATING ROOF | LAWE2 | 94964 | 128898802 | 20850 | 19504 | 147 | 0.81 | 3.48e-05 | 32 | 0 | 18800 | 32 | MM | CRUDE OIL,LAWE | 0 | 32 | 8/27/2023 12:00:00 AM | 8/27/2017 12:00:00 AM | Active
1 | A00001 | R501 | FIXED ROOF | SOUTH | 35976 | 12175465 | 12794 | 12260 | 138 | 0.853 | 3.48e-05 | 0 | 0 | 12000 | 55 | AU | ARJUNA CRUDE | 950 | 0 | 4/30/2020 12:00:00 AM | 4/16/2014 12:00:00 AM | Active
1 | A00002 | R501 | FIXED ROOF | SOUTH | 35972 | 11819391 | 12911 | 12344 | 83 | 0.86 | 3.48e-05 | 65 | 320 | 11500 | 65 | AU | ARJUNA CRUDE | 320 | 65 | 2021-12-02 | 2015-12-02 | Active
1 | A00003 | R501 | FIXED ROOF | SOUTH | 36014 | 12328936 | 13030 | 12299 | -4 | 0.8661 | 3.48e-05 | 0 | 0 | 12100 | 45 | AU | ARJUNA CRUDE | 800 | 0 | 2/28/2019 12:00:00 AM | 2/22/2013 12:00:00 AM | Active
1 | A00004 | R501 | FIXED ROOF | SOUTH | 31963 | 10938799 | 14874 | 14115 | 30 | 0.9193 | 3.48e-05 | 0 | 0 | 13800 | 63 | V2 | LSFO V-1250 | 730 | 0 | 2/28/2019 12:00:00 AM | 2/27/2013 12:00:00 AM | Active
1 | A00005 | R501 | FIXED ROOF | SOUTH | 35972 | 12412894 | 12983 | 12420 | 105 | 0.9034 | 3.48e-05 | 0 | 0 | 12200 | 76 | V2 | LSFO V-1250 | 300 | 0 | 4/30/2020 12:00:00 AM | 4/17/2014 12:00:00 AM | Active
1 | A00006 | R501 | FIXED ROOF | SOUTH | 35932 | 12106776 | 12810 | 12416 | 136 | 0.9327 | 3.48e-05 | 0 | 0 | 12000 | 61 | V1 | LSWR V 1250 | 300 | 0 | 2/28/2019 12:00:00 AM | 2/26/2013 12:00:00 AM | Active
1 | A00007 | R501 | FIXED ROOF | SOUTH | 33999 | 11390323 | 14100 | 12795 | -5 | 0.9602 | 3.48e-05 | 0 | 0 | 12500 | 38 | XO | FLUX OIL | 250 | 0 | 3/26/2021 12:00:00 AM | 3/26/2015 12:00:00 AM | Active
1 | A00015 | R501 | FIXED ROOF | SOUTH | 35933 | 12214308 | 12530 | 12340 | 210 | 0.88 | 3.48e-05 | 41 | 340 | 12040 | 41 | M3 | CRUDE OIL,MUDI | 340 | 41 | 2024-10-05 | 2018-10-05 | Active
1 | A00016 | R501 | FIXED ROOF | SOUTH | 35982 | 12333290 | 12989 | 12130 | 83 | 0.8989 | 3.48e-05 | 15 | 300 | 12130 | 15 | V1 | LSWR V 1250 | 300 | 15 | 2019-11-11 | 1984-11-05 | Active
1 | A00017 | R501 | FIXED ROOF | SOUTH | 35981 | 12306134 | 13040 | 12438 | -10 | 0.914 | 3.48e-05 | 0 | 0 | 12100 | 43 | VG | HVGO INTERMEDIA | 600 | 0 | 2/28/2019 12:00:00 AM | 2/18/2013 12:00:00 AM | Active
1 | A00018 | R501 | FIXED ROOF | SOUTH | 31986 | 11089828 | 14751 | 14108 | 14 | 0.8375 | 3.48e-05 | 0 | 400 | 13800 | 31 | HS | SLOP | 400 | 0 | 2021-05-10 | 2015-05-10 | Active
1 | A00019 | R501 | FIXED ROOF | SOUTH | 31976 | 11248655 | 15402 | 14230 | 75 | 0.8361 | 3.48e-05 | 35 | 400 | 14000 | 35 | HS | SLOP | 400 | 35 | 5/13/2022 12:00:00 AM | 5/13/2016 12:00:00 AM | Active
1 | A00020 | R501 | FIXED ROOF | SOUTH | 31971 | 11167403 | 14555 | 13952 | 96 | 0.995 | 3.48e-05 | 0 | 0 | 13900 | 35 | XO | FLUX OIL | 400 | 0 | 5/31/2019 12:00:00 AM | 5/20/2013 12:00:00 AM | Active
1 | A00021 | R501 | FIXED ROOF | SOUTH | 31903 | 11209885 | 14563 | 14257 | 69 | 0.7437 | 3.48e-05 | 0 | 340 | 14000 | 33 | PR | PREMIUM | 340 | 0 | 11/15/2021 12:00:00 AM | 12/15/2015 12:00:00 AM | Active
1 | A00022 | R501 | FIXED ROOF | SOUTH | 31928 | 11135840 | 14549 | 14170 | 8 | 0.75 | 3.48e-05 | 32 | 300 | 13900 | 32 | PT | PERTAMAX | 300 | 32 | 2022-04-08 | 2016-04-08 | Active
1 | A00023 | R501 | FIXED ROOF | SOUTH | 31936 | 11104547 | 14477 | 14147 | 80 | 0.7437 | 3.48e-05 | 0 | 0 | 13900 | 33 | PR | PREMIUM | 440 | 0 | 3/26/2021 12:00:00 AM | 3/26/2015 12:00:00 AM | Active
1 | A00024 | R501 | FIXED ROOF | SOUTH | 31958 | 11643903 | 15223 | 14894 | 65 | 0.74 | 3.48e-05 | 37 | 440 | 14500 | 37 | PT | PERTAMAX | 440 | 37 | 2024-08-05 | 2018-08-05 | Active
1 | A00025 | R501 | FIXED ROOF | SOUTH | 31975 | 11160973 | 14932 | 14082 | 18 | 0.8137 | 3.48e-05 | 0 | 350 | 13900 | 35 | AV | AVTUR | 350 | 0 | 2023-03-06 | 2017-03-06 | Active
1 | A00026 | R501 | FIXED ROOF | SOUTH | 32054 | 11362991 | 15010 | 14469 | 59 | 0.84 | 3.48e-05 | 100 | 450 | 14100 | 100 | AD | ADO/SOLAR-48 | 450 | 100 | 6/15/2020 12:00:00 AM | 6/15/2014 12:00:00 AM | Active
1 | A00027 | R501 | FIXED ROOF | SOUTH | 31995 | 11058786 | 14535 | 14016 | 149 | 0.82 | 3.48e-05 | 55 | 300 | 13800 | 55 | AV | AVTUR | 300 | 55 | 2022-02-08 | 2016-02-08 | Active
1 | A00028 | R501 | FIXED ROOF | SOUTH | 31995 | 10784924 | 14534 | 13945 | 15 | 0.835 | 3.48e-05 | 34 | 350 | 13400 | 34 | KR | KEROSENE/MINYAK TANAH | 350 | 34 | 9/18/2020 12:00:00 AM | 9/18/2014 12:00:00 AM | Active
1 | A00029 | R501 | FIXED ROOF | SOUTH | 23468 | 4228593 | 10536 | 10035 | 154 | 0.715 | 3.48e-05 | 34 | 430 | 9750 | 34 | EM | EMPTY TANK | 430 | 34 | 9/17/2022 12:00:00 AM | 9/17/2016 12:00:00 AM | Active
1 | A00030 | R501 | FIXED ROOF | SOUTH | 23464 | 4912363 | 12200 | 11572 | 68 | 0.73 | 3.48e-05 | 34 | 400 | 11300 | 34 | PT | PERTAMAX | 400 | 34 | 2/26/2024 12:00:00 AM | 2/26/2018 12:00:00 AM | Active
2 | A00031 | R501 | FLOATING ROOF | SOUTH | 57750 | 39261395 | 17332 | 15955 | 103 | 0.7222 | 3.48e-05 | 45.7 | 380 | 15100 | 29 | PR | PREMIUM | 380 | 45.7 | 2021-05-05 | 2015-05-05 | Active
2 | A00032 | R501 | FLOATING ROOF | SOUTH | 57721 | 35589759 | 17468 | 15934 | -20 | 0.7222 | 3.48e-05 | 29 | 380 | 13650 | 29 | HD | HOMC 92 (HIGH OCTANE MOGAS COMPONENT 92) | 380 | 29 | 2021-11-11 | 2015-11-12 | Active
3 | C2001A | R501 | L P G | NORTH | 23285 | 6607502 | 23000 | 23000 | 230 | 0.5436 | 3.48e-05 | 0 | 0 | 23000 | 26 | PG | LPG | 230 | 0 | 3/31/2017 12:00:00 AM | 3/31/2011 12:00:00 AM | Active
3 | C2001B | R501 | L P G | NORTH | 23276 | 6602774 | 0 | 23276 | 200 | 0.56 | 3.48e-05 | 26 | 200 | 23270 | 26 | PG | LPG | 200 | 26 | 2020-01-09 | 9/30/2014 12:00:00 AM | Active
1 | D20003 | R501 | FIXED ROOF | NORTH | 18991 | 3693199 | 14600 | 13350 | 16 | 0.72 | 3.48e-05 | 55 | 300 | 13000 | 55 | SN | SWEET NAPHTHA | 300 | 55 | 2022-05-08 | 2016-05-08 | Active
1 | D20011 | R501 | FIXED ROOF | NORTH | 26992 | 7402155 | 13677 | 13166 | 0 | 0.8201 | 3.48e-05 | 78 | 400 | 12900 | 78 | HS | SLOP | 400 | 78 | 8/20/2018 12:00:00 AM | 8/20/2016 12:00:00 AM | Active
2 | D20012 | R501 | FLOATING ROOF | NORTH | 27037 | 6947779 | 14815 | 13160 | 0 | 0.9327 | 3.48e-05 | 3.1 | 0 | 12200 | 61 | HS | SLOP | 250 | 3.1 | 3/31/2019 12:00:00 AM | 2/23/2013 12:00:00 AM | Active
1 | D20017 | R501 | FIXED ROOF | NORTH | 23003 | 5406129 | 15025 | 13744 | 106 | 0.8345 | 3.48e-05 | 55 | 300 | 13000 | 55 | EM | EMPTY TANK | 300 | 55 | 2023-07-11 | 2017-07-12 | Active
1 | D20018 | R501 | FIXED ROOF | NORTH | 18602 | 3819194 | 16087 | 14562 | -80 | 0.8345 | 3.48e-05 | 55 | 250 | 14000 | 55 | FL | FLUSHING OIL | 250 | 55 | 2023-07-11 | 2017-07-12 | Active
2 | D2001A | R501 | FLOATING ROOF | NORTH | 40716 | 16381247 | 16458 | 14936 | 160 | 0.8218 | 3.48e-05 | 0 | 160 | 12600 | 55 | MM | CRUDE OIL,LAWE | 160 | 0 | 11/24/2022 12:00:00 AM | 11/24/2016 12:00:00 AM | Active
2 | D2001B | R501 | FLOATING ROOF | NORTH | 40720 | 15680760 | 16452 | 14973 | 156 | 0.822 | 3.48e-05 | 9.8 | 0 | 12100 | 32 | MM | CRUDE OIL,LAWE | 330 | 9.8 | 3/31/2019 12:00:00 AM | 2/25/2013 12:00:00 AM | Active
1 | D2004A | R501 | FIXED ROOF | NORTH | 52158 | 33767716 | 18568 | 16053 | -37 | 0.866 | 3.48e-05 | 0 | 300 | 15700 | 69 | VG | HVGO INTERMEDIA | 300 | 0 | 2019-05-08 | 1998-05-05 | Active
1 | D2004B | R501 | FIXED ROOF | NORTH | 52196 | 33817821 | 18610 | 15940 | 5 | 0.8762 | 3.48e-05 | 0 | 0 | 15700 | 70 | VG | HVGO INTERMEDIA | 150 | 0 | 8/28/2019 12:00:00 AM | 2/28/1997 12:00:00 AM | Active
1 | D2004C | R501 | FIXED ROOF | NORTH | 51212 | 33823598 | 18701 | 15917 | -39 | 0.8653 | 3.48e-05 | 0 | 0 | 15670 | 66 | VG | HVGO INTERMEDIA | 150 | 0 | 8/19/2019 12:00:00 AM | 6/19/1995 12:00:00 AM | Active
2 | D2005A | R501 | FLOATING ROOF | NORTH | 57729 | 39233120 | 17570 | 16001 | 199 | 0.82 | 3.48e-05 | 58 | 500 | 15000 | 58 | FR | FULL RANGE NAPHTHA | 500 | 58 | 9/19/2022 12:00:00 AM | 9/19/2016 12:00:00 AM | Active
2 | D2005B | R501 | FLOATING ROOF | NORTH | 57719 | 38054220 | 17453 | 15972 | 161 | 0.82 | 3.48e-05 | 55 | 350 | 14600 | 55 | KR | KEROSENE/MINYAK TANAH | 350 | 55 | 9/22/2022 12:00:00 AM | 9/22/2016 12:00:00 AM | Active
2 | D2005C | R501 | FLOATING ROOF | NORTH | 57753 | 38420399 | 17443 | 15971 | 167 | 0.81 | 3.48e-05 | 35.2 | 450 | 14600 | 40 | EM | EMPTY TANK | 450 | 35.2 | 6/19/2020 12:00:00 AM | 6/19/2014 12:00:00 AM | Active
2 | D2005D | R501 | FLOATING ROOF | NORTH | 57728 | 38342002 | 17458 | 15997 | 198 | 0.823 | 3.48e-05 | 34.1 | 0 | 14600 | 33 | KR | KEROSENE/MINYAK TANAH | 500 | 34.1 | 4/30/2020 12:00:00 AM | 4/22/2014 12:00:00 AM | Active
2 | D2006A | R501 | FLOATING ROOF | NORTH | 45229 | 24153994 | 17565 | 15980 | 160 | 0.7231 | 3.48e-05 | 28.7 | 0 | 15000 | 36 | FR | FULL RANGE NAPHTHA | 450 | 28.7 | 1/31/2020 12:00:00 AM | 1/20/2014 12:00:00 AM | Active
2 | D2006B | R501 | FLOATING ROOF | NORTH | 45169 | 23976988 | 17445 | 16064 | 179 | 0.8734 | 3.48e-05 | 19.7 | 430 | 15000 | 35 | PF | REFORMATE | 430 | 19.7 | 2020-06-11 | 2014-06-11 | Active
1 | D2007A | R501 | FIXED ROOF | NORTH | 57698 | 41340493 | 16320 | 15970 | 155 | 0.86 | 3.48e-05 | 55 | 450 | 15700 | 55 | AD | ADO/SOLAR-48 | 450 | 55 | 11/13/2023 12:00:00 AM | 12/13/2017 12:00:00 AM | Active
1 | D2007B | R501 | FIXED ROOF | NORTH | 57709 | 41288830 | 16467 | 15980 | 5 | 0.86 | 3.48e-05 | 55 | 450 | 15700 | 55 | AD | ADO/SOLAR-48 | 450 | 55 | 9/28/2022 12:00:00 AM | 9/28/2016 12:00:00 AM | Active
1 | D2007C | R501 | FIXED ROOF | NORTH | 57725 | 40855498 | 16464 | 15992 | 24 | 0.823 | 3.48e-05 | 0 | 0 | 15500 | 35 | AD | ADO/SOLAR-48 | 570 | 0 | 4/30/2020 12:00:00 AM | 4/24/2014 12:00:00 AM | Active
2 | D2013A | R501 | FLOATING ROOF | NORTH | 24494 | 6037403 | 15390 | 13978 | 210 | 0.8065 | 3.48e-05 | 8.6 | 500 | 12900 | 38 | PF | REFORMATE | 500 | 8.6 | 2021-02-10 | 2015-02-10 | Active
2 | D2013B | R501 | FLOATING ROOF | NORTH | 24494 | 6034000 | 15145 | 13988 | 449 | 0.8 | 3.48e-05 | 8.8 | 0 | 12900 | 34 | PF | REFORMATE | 450 | 8.8 | 9/30/2019 12:00:00 AM | 9/26/2013 12:00:00 AM | Active
2 | D2015A | R501 | FLOATING ROOF | NORTH | 18993 | 3396582 | 14805 | 13292 | 96 | 0.7098 | 3.48e-05 | 7.7 | 310 | 12200 | 30 | HT | NAPHTHA INTERMEDIA | 310 | 7.7 | 5/30/2023 12:00:00 AM | 5/30/2017 12:00:00 AM | Active
2 | D2015B | R501 | FLOATING ROOF | NORTH | 18998 | 3411357 | 14752 | 13290 | 168 | 0.7521 | 3.48e-05 | 7.4 | 0 | 12200 | 54 | HT | NAPHTHA INTERMEDIA | 450 | 7.4 | 5/31/2019 12:00:00 AM | 8/21/2013 12:00:00 AM | Active
1 | E00001 | R501 | FIXED ROOF | SOUTH | 15821 | 2346629 | 12198 | 12154 | 11 | 1 | 3.48e-05 | 0 | 360 | 11900 | 25 | EM | EMPTY TANK | 360 | 0 | 2021-01-10 | 2015-01-10 | Active
1 | E00002 | R501 | FIXED ROOF | SOUTH | 15849 | 2329455 | 12645 | 12183 | 50 | 0.89 | 3.48e-05 | 0 | 310 | 11800 | 33 | IF | INDUSTRIAL FUEL OIL / IFO | 310 | 0 | 2/26/2022 12:00:00 AM | 2/26/2016 12:00:00 AM | Active
1 | E00003 | R501 | FIXED ROOF | SOUTH | 15835 | 2330674 | 12201 | 12152 | -3 | 0.8492 | 3.48e-05 | 0 | 300 | 11900 | 36 | ID | INDUSTRIAL DIESEL OILS (IDO),BULK | 300 | 0 | 5/31/2023 12:00:00 AM | 5/31/2017 12:00:00 AM | Active
1 | E00004 | R501 | FIXED ROOF | SOUTH | 15819 | 2251503 | 12324 | 12140 | 0 | 0.9645 | 3.48e-05 | 0 | 400 | 11500 | 35 | IF | INDUSTRIAL FUEL OIL / IFO | 400 | 0 | 2023-01-06 | 2017-01-06 | Active
1 | E00005 | R501 | FIXED ROOF | SOUTH | 15823 | 2335057 | 12360 | 12161 | 15 | 0.8492 | 3.48e-05 | 0 | 50 | 13900 | 36 | ID | INDUSTRIAL DIESEL OILS (IDO),BULK | 50 | 0 | 2023-02-06 | 2017-02-06 | Active
1 | E00006 | R501 | FIXED ROOF | SOUTH | 15834 | 2317908 | 12305 | 12158 | -22 | 0.9542 | 3.48e-05 | 0 | 400 | 11800 | 35 | IF | INDUSTRIAL FUEL OIL / IFO | 400 | 0 | 2021-05-05 | 2015-05-05 | Active
1 | E00007 | R501 | FIXED ROOF | SOUTH | 15849 | 2320258 | 12346 | 12180 | 30 | 0.8438 | 3.48e-05 | 0 | 400 | 11800 | 46 | AD | ADO/SOLAR-48 | 400 | 0 | 2021-07-10 | 2015-07-10 | Active
1 | G00001 | R501 | FIXED ROOF | SOUTH | 7327 | 240323 | 6292 | 5942 | 4 | 0.8498 | 3.48e-05 | 0 | 0 | 5700 | 29 | EM | EMPTY TANK | 0 | 0 | 6/30/2001 12:00:00 AM | 2015-07-12 | Non Active
1 | K00001 | R501 | FIXED ROOF | DISWAX | 15030 | 1508485 | 8666 | 8628 | 14 | 0.789 | 3.48e-05 | 0 | 0 | 8500 | 60 | EM | EMPTY TANK | 40 | 0 | 2/28/2017 12:00:00 AM | 2011-12-02 | Non Active
1 | K00002 | R501 | FIXED ROOF | DISWAX | 15235 | 1626784 | 8771 | 8921 | 279 | 0.8126 | 3.48e-05 | 0 | 0 | 8920 | 42 | EM | EMPTY TANK | 40 | 0 | 9/30/2019 12:00:00 AM | 9/20/2013 12:00:00 AM | Non Active
1 | K00003 | R501 | FIXED ROOF | DISWAX | 15248 | 1643558 | 9469 | 9240 | -18 | 0.8455 | 3.48e-05 | 0 | 0 | 9000 | 50 | EM | EMPTY TANK | 250 | 0 | 2/29/2016 12:00:00 AM | 2/24/2010 12:00:00 AM | Non Active
1 | K00004 | R501 | FIXED ROOF | DISWAX | 5001 | 53021 | 3160 | 3150 | -20 | 0.8455 | 3.48e-05 | 0 | 0 | 3130 | 75 | EM | EMPTY TANK | 0 | 0 | 8/31/2002 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00005 | R501 | FIXED ROOF | DISWAX | 4999 | 57095 | 3050 | 3000 | 50 | 0.8455 | 3.48e-05 | 0 | 0 | 2900 | 75 | EM | EMPTY TANK | 0 | 0 | 8/31/2002 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00006 | R501 | FIXED ROOF | DISWAX | 5003 | 58080 | 3065 | 3006 | 20 | 0.8455 | 3.48e-05 | 0 | 0 | 2950 | 60 | EM | EMPTY TANK | 0 | 0 | 12/31/2002 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00007 | R501 | FIXED ROOF | DISWAX | 4796 | 106669 | 6093 | 5988 | -5 | 0.8455 | 3.48e-05 | 0 | 0 | 5900 | 30 | EM | EMPTY TANK | 0 | 0 | 4/30/1998 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00008 | R501 | FIXED ROOF | DISWAX | 4797 | 106547 | 6084 | 5988 | -9 | 0.8455 | 3.48e-05 | 0 | 0 | 5900 | 27 | EM | EMPTY TANK | 0 | 0 | 4/30/1998 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00009 | R501 | FIXED ROOF | DISWAX | 4788 | 102244 | 6087 | 5977 | -25 | 0.8455 | 3.48e-05 | 0 | 0 | 5680 | 30 | EM | EMPTY TANK | 0 | 0 | 2/28/1999 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00010 | R501 | FIXED ROOF | DISWAX | 5000 | 49465 | 3067 | 2986 | -13 | 0.8455 | 3.48e-05 | 0 | 0 | 2540 | 25 | EM | EMPTY TANK | 0 | 0 | 2/28/1999 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00011 | R501 | FIXED ROOF | DISWAX | 4990 | 48911 | 3062 | 2989 | 19 | 0.8455 | 3.48e-05 | 0 | 0 | 2540 | 30 | EM | EMPTY TANK | 0 | 0 | 3/31/1999 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00012 | R501 | FIXED ROOF | DISWAX | 2995 | 19112 | 2777 | 2750 | -3 | 0.8455 | 3.48e-05 | 0 | 0 | 2650 | 41 | EM | EMPTY TANK | 0 | 0 | 9/30/1998 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00013 | R501 | FIXED ROOF | DISWAX | 2996 | 9062 | 1510 | 1500 | 15 | 0.8475 | 3.48e-05 | 0 | 0 | 1290 | 55 | EM | EMPTY TANK | 200 | 0 | 5/31/2016 12:00:00 AM | 2010-04-05 | Non Active
1 | K00014 | R501 | FIXED ROOF | DISWAX | 2992 | 10171 | 1510 | 1490 | 50 | 0.789 | 3.48e-05 | 0 | 0 | 1450 | 55 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/16/2011 12:00:00 AM | Non Active
1 | K00015 | R501 | FIXED ROOF | DISWAX | 2994 | 10183 | 1537 | 1500 | 3 | 0.799 | 3.48e-05 | 0 | 0 | 1450 | 85 | EM | EMPTY TANK | 200 | 0 | 2/28/2017 12:00:00 AM | 2/17/2011 12:00:00 AM | Non Active
1 | K00016 | R501 | FIXED ROOF | DISWAX | 2999 | 9152 | 1522 | 1500 | 6 | 0.8475 | 3.48e-05 | 0 | 0 | 1300 | 55 | EM | EMPTY TANK | 200 | 0 | 5/31/2016 12:00:00 AM | 2010-04-05 | Non Active
1 | K00017 | R501 | FIXED ROOF | DISWAX | 3001 | 9094 | 1530 | 1490 | -6 | 0.8444 | 3.48e-05 | 0 | 0 | 1290 | 60 | EM | EMPTY TANK | 200 | 0 | 5/31/2016 12:00:00 AM | 2010-04-05 | Non Active
1 | K00018 | R501 | FIXED ROOF | DISWAX | 2996 | 8867 | 1510 | 1465 | 2 | 0.814 | 3.48e-05 | 0 | 0 | 1260 | 65 | EM | EMPTY TANK | 200 | 0 | 5/31/2016 12:00:00 AM | 2010-04-05 | Non Active
1 | K00019 | R501 | FIXED ROOF | DISWAX | 3000 | 10218 | 1538 | 1502 | 5 | 0.789 | 3.48e-05 | 0 | 0 | 1450 | 55 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/18/2011 12:00:00 AM | Non Active
1 | K00020 | R501 | FIXED ROOF | DISWAX | 3002 | 9169 | 1535 | 1500 | 3 | 0.8336 | 3.48e-05 | 0 | 0 | 1300 | 70 | EM | EMPTY TANK | 200 | 0 | 7/31/2019 12:00:00 AM | 7/25/2013 12:00:00 AM | Non Active
1 | K00021 | R501 | FIXED ROOF | DISWAX | 3001 | 10227 | 1385 | 1495 | 155 | 0.822 | 3.48e-05 | 0 | 0 | 1450 | 70 | EM | EMPTY TANK | 100 | 0 | 7/31/2019 12:00:00 AM | 7/25/2013 12:00:00 AM | Non Active
1 | K00022 | R501 | FIXED ROOF | DISWAX | 5999 | 121462 | 7243 | 7208 | 13 | 0.8455 | 3.48e-05 | 0 | 0 | 7000 | 60 | EM | EMPTY TANK | 0 | 0 | 11/30/2014 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00023 | R501 | FIXED ROOF | DISWAX | 6015 | 198638 | 7291 | 7230 | 45 | 0.79 | 3.48e-05 | 0 | 0 | 7000 | 55 | EM | EMPTY TANK | 0 | 0 | 11/30/2015 12:00:00 AM | 2013-12-11 | Non Active
1 | K00024 | R501 | FIXED ROOF | DISWAX | 5995 | 197683 | 7262 | 7212 | 16 | 0.792 | 3.48e-05 | 47 | 300 | 7000 | 47 | EM | EMPTY TANK | 300 | 47 | 2020-06-11 | 2014-06-11 | Non Active
1 | K00025 | R501 | FIXED ROOF | DISWAX | 6002 | 198025 | 7273 | 7192 | 14 | 0.8734 | 3.48e-05 | 40 | 260 | 7000 | 40 | EM | EMPTY TANK | 260 | 40 | 2020-06-11 | 2014-06-11 | Non Active
1 | K00026 | R501 | FIXED ROOF | DISWAX | 6004 | 195361 | 7235 | 7192 | 4 | 0.8734 | 3.48e-05 | 35 | 260 | 6900 | 35 | EM | EMPTY TANK | 260 | 35 | 2020-06-11 | 2014-06-11 | Non Active
1 | K00027 | R501 | FIXED ROOF | DISWAX | 6002 | 198179 | 7310 | 7248 | 2 | 0.8734 | 3.48e-05 | 35 | 260 | 7000 | 35 | EM | EMPTY TANK | 260 | 35 | 2020-07-11 | 2014-07-11 | Non Active
1 | K00028 | R501 | FIXED ROOF | DISWAX | 6001 | 198104 | 7275 | 7240 | 0 | 0.856 | 3.48e-05 | 0 | 0 | 7000 | 70 | EM | EMPTY TANK | 0 | 0 | 7/31/2019 12:00:00 AM | 7/27/2013 12:00:00 AM | Non Active
1 | K00029 | R501 | FIXED ROOF | DISWAX | 6005 | 197925 | 7268 | 7240 | -13 | 0.811 | 3.48e-05 | 0 | 0 | 7000 | 60 | EM | EMPTY TANK | 0 | 0 | 2/28/2019 12:00:00 AM | 2/28/2013 12:00:00 AM | Non Active
1 | K00030 | R501 | FIXED ROOF | DISWAX | 6005 | 197981 | 7272 | 7235 | -12 | 0.811 | 3.48e-05 | 0 | 0 | 7000 | 60 | EM | EMPTY TANK | 0 | 0 | 2/28/2019 12:00:00 AM | 2/28/2013 12:00:00 AM | Non Active
1 | K00031 | R501 | FIXED ROOF | DISWAX | 6010 | 198408 | 7243 | 7160 | -12 | 0.835 | 3.48e-05 | 70 | 260 | 7000 | 70 | EM | EMPTY TANK | 260 | 70 | 9/16/2020 12:00:00 AM | 9/16/2014 12:00:00 AM | Non Active
1 | K00032 | R501 | FIXED ROOF | DISWAX | 5993 | 197211 | 7238 | 7232 | 118 | 0.792 | 3.48e-05 | 47 | 260 | 7000 | 47 | EM | EMPTY TANK | 260 | 47 | 2020-07-11 | 2014-07-11 | Non Active
1 | K00033 | R501 | FIXED ROOF | DISWAX | 5995 | 197677 | 7263 | 7217 | 0 | 0.8479 | 3.48e-05 | 0 | 0 | 7000 | 60 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/15/2013 12:00:00 AM | Non Active
1 | K00034 | R501 | FIXED ROOF | DISWAX | 5991 | 197285 | 7297 | 7257 | 81 | 0.7887 | 3.48e-05 | 0 | 0 | 7000 | 55 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/25/2011 12:00:00 AM | Non Active
1 | K00035 | R501 | FIXED ROOF | DISWAX | 5994 | 197573 | 7286 | 7241 | 10 | 0.788 | 3.48e-05 | 0 | 0 | 7000 | 55 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/24/2011 12:00:00 AM | Non Active
1 | K00036 | R501 | FIXED ROOF | DISWAX | 5996 | 197627 | 7214 | 7190 | -12 | 0.799 | 3.48e-05 | 0 | 0 | 7000 | 50 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/23/2011 12:00:00 AM | Non Active
1 | K00037 | R501 | FIXED ROOF | DISWAX | 6001 | 197957 | 7210 | 7195 | -15 | 0.8538 | 3.48e-05 | 0 | 0 | 7000 | 70 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/20/2013 12:00:00 AM | Non Active
1 | K00038 | R501 | FIXED ROOF | DISWAX | 6000 | 192120 | 7093 | 7085 | 17 | 0.792 | 3.48e-05 | 47 | 360 | 6800 | 47 | EM | EMPTY TANK | 360 | 47 | 2020-07-11 | 2014-07-11 | Non Active
1 | K00039 | R501 | FIXED ROOF | DISWAX | 6029 | 193931 | 7109 | 7095 | 9 | 0.792 | 3.48e-05 | 47 | 360 | 6800 | 47 | EM | EMPTY TANK | 360 | 47 | 2020-07-11 | 2014-07-11 | Non Active
1 | K00040 | R501 | FIXED ROOF | DISWAX | 9016 | 542643 | 8769 | 8730 | 10 | 0.798 | 3.48e-05 | 0 | 0 | 8500 | 55 | EM | EMPTY TANK | 0 | 0 | 2/28/2017 12:00:00 AM | 2/21/2011 12:00:00 AM | Non Active
1 | K00041 | R501 | FIXED ROOF | DISWAX | 9006 | 538226 | 8794 | 8715 | -12 | 0.8538 | 3.48e-05 | 0 | 0 | 8450 | 70 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/16/2013 12:00:00 AM | Non Active
1 | K00042 | R501 | FIXED ROOF | DISWAX | 9004 | 541171 | 8758 | 8691 | -15 | 0.8538 | 3.48e-05 | 0 | 0 | 8500 | 70 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/16/2013 12:00:00 AM | Non Active
1 | K00043 | R501 | FIXED ROOF | DISWAX | 8999 | 536068 | 8834 | 8717 | -37 | 0.8538 | 3.48e-05 | 0 | 0 | 8430 | 70 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/17/2013 12:00:00 AM | Non Active
1 | K00044 | R501 | FIXED ROOF | DISWAX | 8002 | 236397 | 5101 | 4965 | -3 | 0.8213 | 3.48e-05 | 0 | 0 | 4700 | 50 | EM | EMPTY TANK | 0 | 0 | 4/30/2019 12:00:00 AM | 4/16/2013 12:00:00 AM | Non Active
1 | K00045 | R501 | FIXED ROOF | DISWAX | 7993 | 233394 | 5019 | 4907 | 4 | 0.792 | 3.48e-05 | 47 | 240 | 4650 | 47 | EM | EMPTY TANK | 240 | 47 | 2020-08-11 | 2014-08-11 | Non Active
1 | K00046 | R501 | FIXED ROOF | DISWAX | 8969 | 537038 | 8768 | 8715 | -4 | 0.8046 | 3.48e-05 | 0 | 0 | 8500 | 50 | EM | EMPTY TANK | 40 | 0 | 4/30/2019 12:00:00 AM | 4/17/2013 12:00:00 AM | Non Active
1 | K00047 | R501 | FIXED ROOF | DISWAX | 7318 | 305709 | 7576 | 7515 | -1 | 0.8147 | 3.48e-05 | 0 | 0 | 7250 | 50 | EM | EMPTY TANK | 40 | 0 | 4/30/2019 12:00:00 AM | 4/15/2013 12:00:00 AM | Non Active
1 | K00048 | R501 | FIXED ROOF | DISWAX | 6001 | 197942 | 7155 | 7125 | 0 | 0.8191 | 3.48e-05 | 0 | 0 | 7000 | 70 | EM | EMPTY TANK | 40 | 0 | 7/31/2019 12:00:00 AM | 7/28/2013 12:00:00 AM | Non Active
1 | K00049 | R501 | FIXED ROOF | DISWAX | 5991 | 197294 | 7152 | 7145 | 11 | 0.7974 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 40 | 0 | 5/31/2019 12:00:00 AM | 5/15/2013 12:00:00 AM | Non Active
1 | K00050 | R501 | FIXED ROOF | DISWAX | 5998 | 197665 | 7266 | 7255 | 5 | 0.812 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 40 | 0 | 5/31/2019 12:00:00 AM | 5/15/2013 12:00:00 AM | Non Active
1 | K00051 | R501 | FIXED ROOF | DISWAX | 5991 | 197129 | 7146 | 7146 | 25 | 0.8233 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 350 | 0 | 5/31/2019 12:00:00 AM | 5/18/2013 12:00:00 AM | Non Active
1 | K00052 | R501 | FIXED ROOF | DISWAX | 5987 | 197208 | 7144 | 7133 | 8 | 0.8132 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 40 | 0 | 5/31/2019 12:00:00 AM | 5/17/2013 12:00:00 AM | Non Active
1 | K00053 | R501 | FIXED ROOF | DISWAX | 5988 | 197251 | 7216 | 7183 | -13 | 0.8146 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 40 | 0 | 5/31/2019 12:00:00 AM | 5/18/2013 12:00:00 AM | Non Active
1 | K00054 | R501 | FIXED ROOF | DISWAX | 5998 | 197779 | 7237 | 7219 | 11 | 0.792 | 3.48e-05 | 47 | 350 | 7000 | 47 | EM | EMPTY TANK | 350 | 47 | 2020-08-11 | 2014-08-11 | Non Active
1 | K00055 | R501 | FIXED ROOF | DISWAX | 5995 | 197684 | 7218 | 7215 | 4 | 0.792 | 3.48e-05 | 47 | 350 | 7000 | 47 | EM | EMPTY TANK | 350 | 47 | 2020-08-11 | 2014-08-11 | Non Active
1 | K00056 | R501 | FIXED ROOF | DISWAX | 6000 | 198109 | 7274 | 7250 | 2 | 0.792 | 3.48e-05 | 47 | 350 | 7000 | 47 | EM | EMPTY TANK | 350 | 47 | 2020-08-11 | 2014-08-11 | Non Active
1 | K00057 | R501 | FIXED ROOF | DISWAX | 5999 | 197979 | 7223 | 9195 | 3 | 0.8734 | 3.48e-05 | 35 | 350 | 7000 | 35 | EM | EMPTY TANK | 350 | 35 | 2020-10-11 | 2014-10-11 | Non Active
1 | K00058 | R501 | FIXED ROOF | DISWAX | 5997 | 197527 | 7246 | 7175 | 52 | 0.792 | 3.48e-05 | 47 | 350 | 7000 | 47 | EM | EMPTY TANK | 350 | 47 | 2020-10-11 | 2014-10-11 | Non Active
1 | K00059 | R501 | FIXED ROOF | DISWAX | 6008 | 198433 | 7245 | 7200 | 5 | 0.812 | 3.48e-05 | 0 | 0 | 7000 | 45 | EM | EMPTY TANK | 40 | 0 | 7/31/2019 12:00:00 AM | 7/28/2013 12:00:00 AM | Non Active
1 | K00060 | R501 | FIXED ROOF | DISWAX | 6006 | 198264 | 7208 | 7207 | 0 | 0.811 | 3.48e-05 | 70 | 350 | 7000 | 70 | EM | EMPTY TANK | 350 | 70 | 9/16/2020 12:00:00 AM | 9/16/2014 12:00:00 AM | Non Active
1 | K00061 | R501 | FIXED ROOF | DISWAX | 6004 | 198185 | 7240 | 7188 | 0 | 0.811 | 3.48e-05 | 70 | 350 | 7000 | 70 | EM | EMPTY TANK | 350 | 70 | 9/16/2020 12:00:00 AM | 9/16/2014 12:00:00 AM | Non Active
1 | K00062 | R501 | FIXED ROOF | DISWAX | 6008 | 198342 | 7252 | 7205 | 0 | 0.811 | 3.48e-05 | 70 | 350 | 7000 | 70 | EM | EMPTY TANK | 350 | 70 | 9/16/2020 12:00:00 AM | 9/16/2014 12:00:00 AM | Non Active
1 | K00063 | R501 | FIXED ROOF | DISWAX | 3001 | 19578 | 3074 | 2995 | 52 | 0.8182 | 3.48e-05 | 0 | 0 | 2700 | 60 | EM | EMPTY TANK | 40 | 0 | 7/31/2019 12:00:00 AM | 7/25/2013 12:00:00 AM | Non Active
1 | K00064 | R501 | FIXED ROOF | DISWAX | 3010 | 18855 | 2921 | 2919 | 11 | 0.8182 | 3.48e-05 | 0 | 0 | 2650 | 60 | EM | EMPTY TANK | 40 | 0 | 7/31/2019 12:00:00 AM | 7/26/2013 12:00:00 AM | Non Active
1 | K00065 | R501 | FIXED ROOF | DISWAX | 3047 | 24066 | 3602 | 3560 | 8 | 0.8182 | 3.48e-05 | 0 | 0 | 3300 | 60 | EM | EMPTY TANK | 40 | 0 | 7/31/2019 12:00:00 AM | 7/26/2013 12:00:00 AM | Non Active
1 | K00066 | R501 | FIXED ROOF | DISWAX | 3411 | 24732 | 3178 | 3100 | -24 | 0.8156 | 3.48e-05 | 0 | 0 | 2700 | 70 | EM | EMPTY TANK | 290 | 0 | 7/31/2019 12:00:00 AM | 7/26/2013 12:00:00 AM | Non Active
1 | K00067 | R501 | FIXED ROOF | DISWAX | 3464 | 26257 | 3010 | 3050 | 60 | 0.8156 | 3.48e-05 | 0 | 0 | 2800 | 70 | EM | EMPTY TANK | 0 | 0 | 7/31/2019 12:00:00 AM | 7/26/2013 12:00:00 AM | Non Active
1 | K00068 | R501 | FIXED ROOF | DISWAX | 4797 | 102974 | 6034 | 6006 | 16 | 0.8113 | 3.48e-05 | 0 | 0 | 5700 | 45 | EM | EMPTY TANK | 40 | 0 | 5/31/2019 12:00:00 AM | 2013-08-05 | Non Active
1 | K00069 | R501 | FIXED ROOF | DISWAX | 4806 | 104283 | 5980 | 6026 | 144 | 0.8034 | 3.48e-05 | 0 | 0 | 5750 | 45 | EM | EMPTY TANK | 260 | 0 | 5/31/2019 12:00:00 AM | 5/15/2013 12:00:00 AM | Non Active
1 | K00070 | R501 | FIXED ROOF | DISWAX | 4804 | 103199 | 6208 | 5975 | 30 | 0.8183 | 3.48e-05 | 0 | 0 | 5700 | 63 | EM | EMPTY TANK | 260 | 0 | 2/29/2016 12:00:00 AM | 2/25/2010 12:00:00 AM | Non Active
1 | K00071 | R501 | FIXED ROOF | DISWAX | 2998 | 19011 | 2923 | 2824 | 5 | 0.778 | 3.48e-05 | 0 | 0 | 2700 | 60 | EM | EMPTY TANK | 0 | 0 | 11/30/2015 12:00:00 AM | 2013-12-11 | Non Active
1 | K00072 | R501 | FIXED ROOF | DISWAX | 2986 | 18874 | 2954 | 2784 | 21 | 0.778 | 3.48e-05 | 0 | 0 | 2700 | 60 | EM | EMPTY TANK | 0 | 0 | 11/30/2015 12:00:00 AM | 2013-12-11 | Non Active
1 | K00073 | R501 | FIXED ROOF | DISWAX | 2998 | 18740 | 2952 | 2791 | -31 | 0.8 | 3.48e-05 | 0 | 0 | 2660 | 65 | EM | EMPTY TANK | 0 | 0 | 11/30/1998 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00074 | R501 | FIXED ROOF | DISWAX | 2999 | 18752 | 2943 | 2779 | -6 | 0.8 | 3.48e-05 | 0 | 0 | 2660 | 65 | EM | EMPTY TANK | 0 | 0 | 11/30/1998 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | K00075 | R501 | FIXED ROOF | DISWAX | 9001 | 559603 | 9315 | 9102 | 19 | 0.792 | 3.48e-05 | 47 | 300 | 8800 | 47 | EM | EMPTY TANK | 300 | 47 | 2020-10-11 | 2014-10-11 | Non Active
1 | K00076 | R501 | FIXED ROOF | DISWAX | 9007 | 567213 | 9300 | 9125 | 5 | 0.792 | 3.48e-05 | 47 | 300 | 8900 | 47 | EM | EMPTY TANK | 300 | 47 | 2020-10-11 | 2014-10-11 | Non Active
1 | K00077 | R501 | FIXED ROOF | DISWAX | 8998 | 559910 | 9300 | 9102 | -4 | 0.792 | 3.48e-05 | 47 | 450 | 8800 | 47 | EM | EMPTY TANK | 450 | 47 | 2020-10-11 | 2014-10-11 | Non Active
1 | K00078 | R501 | FIXED ROOF | DISWAX | 9004 | 572776 | 9313 | 9125 | 48 | 0.776 | 3.48e-05 | 0 | 0 | 9000 | 60 | EM | EMPTY TANK | 40 | 0 | 10/31/2015 12:00:00 AM | 10/27/2013 12:00:00 AM | Non Active
1 | K00079 | R501 | FIXED ROOF | DISWAX | 9003 | 566595 | 9285 | 9105 | 5 | 0.792 | 3.48e-05 | 47 | 300 | 8900 | 47 | EM | EMPTY TANK | 300 | 47 | 2020-11-11 | 2014-11-11 | Non Active
1 | K00080 | R501 | FIXED ROOF | DISWAX | 9005 | 560939 | 9275 | 9122 | 3 | 0.8734 | 3.48e-05 | 35 | 450 | 8800 | 35 | EM | EMPTY TANK | 450 | 35 | 2020-11-11 | 2014-11-11 | Non Active
1 | K00081 | R501 | FIXED ROOF | DISWAX | 6198 | 232201 | 8168 | 7932 | 4 | 0.789 | 3.48e-05 | 0 | 0 | 7700 | 50 | EM | EMPTY TANK | 270 | 0 | 2/29/2016 12:00:00 AM | 2/24/2010 12:00:00 AM | Non Active
1 | K00082 | R501 | FIXED ROOF | DISWAX | 6214 | 230676 | 8105 | 7880 | 0 | 0.8734 | 3.48e-05 | 35 | 270 | 7600 | 35 | EM | EMPTY TANK | 270 | 35 | 2020-11-11 | 2014-11-11 | Non Active
1 | K00083 | R501 | FIXED ROOF | DISWAX | 5978 | 213583 | 8087 | 7844 | 1 | 0.8734 | 3.48e-05 | 35 | 270 | 7600 | 35 | EM | EMPTY TANK | 270 | 35 | 2020-12-11 | 2014-12-11 | Non Active
1 | K00084 | R501 | FIXED ROOF | DISWAX | 5973 | 209998 | 7929 | 7725 | 3 | 0.8734 | 3.48e-05 | 35 | 270 | 7500 | 35 | EM | EMPTY TANK | 270 | 35 | 2020-12-11 | 2014-12-11 | Non Active
1 | K00085 | R501 | FIXED ROOF | DISWAX | 3198 | 22486 | 3137 | 3018 | 37 | 0.784 | 3.48e-05 | 0 | 0 | 2800 | 58 | EM | EMPTY TANK | 0 | 0 | 11/30/2015 12:00:00 AM | 2013-12-11 | Non Active
1 | K0019A | R501 | FIXED ROOF | DISWAX | 2999 | 10240 | 1617 | 1505 | -2 | 0.8167 | 3.48e-05 | 0 | 0 | 1450 | 70 | EM | EMPTY TANK | 0 | 0 | 7/31/2019 12:00:00 AM | 7/25/2013 12:00:00 AM | Non Active
1 | K0020A | R501 | FIXED ROOF | DISWAX | 3000 | 8945 | 1568 | 1477 | 12 | 0.8444 | 3.48e-05 | 0 | 0 | 1270 | 60 | EM | EMPTY TANK | 200 | 0 | 5/31/2016 12:00:00 AM | 2010-04-05 | Non Active
1 | K0021A | R501 | FIXED ROOF | DISWAX | 3003 | 10237 | 1524 | 1502 | 6 | 0.8 | 3.48e-05 | 0 | 0 | 1450 | 65 | EM | EMPTY TANK | 100 | 0 | 4/30/2009 12:00:00 AM | 2015-07-12 | Non Active
1 | N00001 | R501 | FIXED ROOF | NORTH | 12199 | 1250800 | 11125 | 10962 | 5 | 0.8734 | 3.48e-05 | 35 | 420 | 10700 | 35 | FL | FLUSHING OIL | 420 | 35 | 2020-12-11 | 2014-12-11 | Active
1 | N00002 | R501 | FIXED ROOF | NORTH | 8175 | 293943 | 6145 | 5900 | 48 | 0.8201 | 3.48e-05 | 0 | 0 | 5600 | 32 | S0 | SMOOTH FLUID 05 | 250 | 0 | 1/31/2020 12:00:00 AM | 1/23/2014 12:00:00 AM | Active
1 | N00003 | R501 | FIXED ROOF | NORTH | 7494 | 308833 | 7410 | 7292 | 23 | 0.8422 | 3.48e-05 | 0 | 0 | 7000 | 60 | HS | SLOP | 0 | 0 | 9/30/2019 12:00:00 AM | 9/27/2013 12:00:00 AM | Active
1 | N00004 | R501 | FIXED ROOF | NORTH | 7472 | 307111 | 7688 | 7294 | -20 | 0.84 | 3.48e-05 | 35 | 250 | 7000 | 40 | S0 | SMOOTH FLUID 05 | 250 | 35 | 9/16/2020 12:00:00 AM | 9/16/2014 12:00:00 AM | Active
1 | N00005 | R501 | FIXED ROOF | NORTH | 7500 | 315395 | 7940 | 7420 | -5 | 0.8502 | 3.48e-05 | 0 | 0 | 7140 | 40 | EM | EMPTY TANK | 0 | 0 | 10/31/2015 12:00:00 AM | 10/27/2013 12:00:00 AM | Active
1 | N00006 | R501 | FIXED ROOF | NORTH | 12009 | 577691 | 5816 | 5382 | -1 | 0.8424 | 3.48e-05 | 0 | 0 | 5100 | 60 | HS | SLOP | 0 | 0 | 9/30/2019 12:00:00 AM | 9/27/2013 12:00:00 AM | Active
1 | N00024 | R501 | FIXED ROOF | DISWAX | 12193 | 397209 | 3989 | 3542 | 7 | 0.01 | 3.48e-05 | 0 | 0 | 3400 | 40 | EM | EMPTY TANK | 0 | 0 | 6/30/2001 12:00:00 AM | 1/25/2017 12:00:00 AM | Non Active
1 | O00001 | R501 | FIXED ROOF | NORTH | 36012 | 12346893 | 12755 | 12420 | 130 | 0.9191 | 3.48e-05 | 70 | 300 | 12120 | 70 | R6 | SHORT RESIDUE | 300 | 70 | 5/24/2020 12:00:00 AM | 3/24/1999 12:00:00 AM | Active
1 | O00002 | R501 | FIXED ROOF | NORTH | 35996 | 10512759 | 11301 | 10647 | 7 | 0.9 | 3.48e-05 | 35 | 750 | 10300 | 35 | R6 | SHORT RESIDUE | 750 | 35 | 9/30/2020 12:00:00 AM | 9/30/2014 12:00:00 AM | Active
1 | O00003 | R501 | FIXED ROOF | NORTH | 35994 | 12131038 | 12536 | 11850 | 5 | 0.9878 | 3.48e-05 | 73 | 700 | 11850 | 73 | OR | REFINERY FUEL OIL | 700 | 73 | 11/29/2019 12:00:00 AM | 1/29/2000 12:00:00 AM | Active
1 | O00004 | R501 | FIXED ROOF | NORTH | 35992 | 12055765 | 12345 | 12125 | 104 | 0.91 | 3.48e-05 | 120 | 700 | 11850 | 120 | EM | EMPTY TANK | 700 | 120 | 9/30/2020 12:00:00 AM | 9/30/2014 12:00:00 AM | Active
1 | O00006 | R501 | FIXED ROOF | NORTH | 36006 | 12350271 | 12645 | 12000 | 65 | 0.9035 | 3.48e-05 | 77 | 300 | 12000 | 77 | R6 | SHORT RESIDUE | 300 | 77 | 11/19/2019 12:00:00 AM | 7/19/2008 12:00:00 AM | Active
1 | O00008 | R501 | FIXED ROOF | NORTH | 35952 | 12237209 | 12990 | 12100 | -20 | 0.94 | 3.48e-05 | 91 | 510 | 12100 | 91 | R6 | SHORT RESIDUE | 510 | 91 | 2019-11-11 | 2010-11-06 | Active
1 | O00009 | R501 | FIXED ROOF | NORTH | 23991 | 4883319 | 11466 | 11128 | 9 | 0.9462 | 3.48e-05 | 52 | 300 | 10800 | 52 | OR | REFINERY FUEL OIL | 300 | 52 | 5/15/2022 12:00:00 AM | 5/15/2016 12:00:00 AM | Active
1 | O00010 | R501 | FIXED ROOF | NORTH | 24374 | 5513903 | 12242 | 12126 | 24 | 0.8482 | 3.48e-05 | 0 | 0 | 11800 | 40 | XO | FLUX OIL | 400 | 0 | 1/31/2019 12:00:00 AM | 2013-11-01 | Active
1 | O00011 | R501 | FIXED ROOF | NORTH | 23994 | 4883215 | 11491 | 11126 | 8 | 0.74 | 3.48e-05 | 85 | 300 | 10800 | 85 | EM | EMPTY TANK | 300 | 85 | 2022-05-08 | 2016-05-08 | Active
1 | O00012 | R501 | FIXED ROOF | NORTH | 24393 | 5512375 | 12329 | 11800 | 120 | 0.8429 | 3.48e-05 | 75 | 750 | 11800 | 75 | OR | REFINERY FUEL OIL | 750 | 75 | 11/14/2019 12:00:00 AM | 6/14/1983 12:00:00 AM | Active
1 | O00013 | R501 | FIXED ROOF | NORTH | 21321 | 4178807 | 12170 | 12143 | 200 | 0.74 | 3.48e-05 | 35 | 750 | 11800 | 35 | XO | FLUX OIL | 750 | 35 | 2022-08-08 | 2016-08-08 | Active
1 | O00014 | R501 | FIXED ROOF | NORTH | 24372 | 5537672 | 0 | 12114 | -20 | 0.8932 | 3.48e-05 | 72 | 750 | 11900 | 72 | OR | REFINERY FUEL OIL | 750 | 72 | 3/20/2020 12:00:00 AM | 11/20/1998 12:00:00 AM | Active
1 | O00015 | R501 | FIXED ROOF | NORTH | 23986 | 5515013 | 13050 | 12396 | -105 | 0.8454 | 3.48e-05 | 0 | 0 | 12100 | 45 | SC | PERTADEX | 350 | 0 | 9/30/2019 12:00:00 AM | 9/22/2013 12:00:00 AM | Active
1 | O00016 | R501 | FIXED ROOF | NORTH | 21347 | 4226122 | 12426 | 12186 | 58 | 0.8438 | 3.48e-05 | 0 | 300 | 11800 | 46 | SC | PERTADEX | 300 | 0 | 11/16/2021 12:00:00 AM | 12/16/2015 12:00:00 AM | Active
1 | O00017 | R501 | FIXED ROOF | NORTH | 23982 | 5468202 | 12954 | 12371 | 9 | 0.74 | 3.48e-05 | 38 | 400 | 12100 | 38 | AD | ADO/SOLAR-48 | 400 | 38 | 2022-09-08 | 2016-09-08 | Active
1 | O00018 | R501 | FIXED ROOF | NORTH | 21309 | 4117941 | 12425 | 12110 | 139 | 0.85 | 3.48e-05 | 90 | 230 | 11500 | 90 | XO | FLUX OIL | 230 | 90 | 2024-06-05 | 2018-06-05 | Active
1 | O00019 | R501 | FIXED ROOF | NORTH | 23990 | 5540763 | 12660 | 12337 | 64 | 0.8604 | 3.48e-05 | 40 | 200 | 12100 | 40 | XO | FLUX OIL | 200 | 40 | 6/19/2020 12:00:00 AM | 6/19/2014 12:00:00 AM | Active
1 | O00020 | R501 | FIXED ROOF | NORTH | 21323 | 4287825 | 12325 | 12104 | -62 | 0.82 | 3.48e-05 | 0 | 0 | 12000 | 45 | S0 | SMOOTH FLUID 05 | 320 | 0 | 5/31/2019 12:00:00 AM | 5/20/2013 12:00:00 AM | Active
1 | O00021 | R501 | FIXED ROOF | NORTH | 23988 | 5476785 | 12812 | 12357 | 174 | 0.85 | 3.48e-05 | 58 | 400 | 12100 | 58 | VG | HVGO INTERMEDIA | 400 | 58 | 9/26/2022 12:00:00 AM | 9/26/2016 12:00:00 AM | Active
1 | O00022 | R501 | FIXED ROOF | NORTH | 21335 | 4243247 | 12323 | 12160 | -24 | 0.838 | 3.48e-05 | 0 | 0 | 11900 | 61 | NF | NET BOT.FRACTIONATOR | 350 | 0 | 1/31/2019 12:00:00 AM | 2013-10-01 | Active
1 | O00023 | R501 | FIXED ROOF | NORTH | 24001 | 5475468 | 12719 | 12400 | -9 | 0.8523 | 3.48e-05 | 0 | 0 | 12100 | 50 | NF | NET BOT.FRACTIONATOR | 300 | 0 | 3/26/2021 12:00:00 AM | 3/26/2015 12:00:00 AM | Active
1 | O00024 | R501 | FIXED ROOF | NORTH | 39009 | 15091117 | 14063 | 12792 | 35 | 0.9 | 3.48e-05 | 52 | 260 | 12500 | 52 | R6 | SHORT RESIDUE | 260 | 52 | 5/14/2022 12:00:00 AM | 5/14/2016 12:00:00 AM | Active
1 | O00025 | R501 | FIXED ROOF | NORTH | 39035 | 15031391 | 14152 | 12861 | -55 | 0.9 | 3.48e-05 | 0 | 280 | 12500 | 79 | R6 | SHORT RESIDUE | 280 | 0 | 2022-03-03 | 2016-03-03 | Active
1 | P00001 | R501 | FIXED ROOF | NORTH | 23985 | 5482836 | 12632 | 12349 | 106 | 0.883 | 3.48e-05 | 39 | 400 | 12100 | 39 | AV | AVTUR | 400 | 39 | 11/13/2023 12:00:00 AM | 12/13/2017 12:00:00 AM | Active
1 | P00002 | R501 | FIXED ROOF | NORTH | 23994 | 5410059 | 13077 | 12195 | 31 | 0.806 | 3.48e-05 | 0 | 0 | 11900 | 41 | AV | AVTUR | 340 | 0 | 2/28/2019 12:00:00 AM | 2/21/2013 12:00:00 AM | Active
2 | P00003 | R501 | FLOATING ROOF | NORTH | 39018 | 14288113 | 15367 | 12808 | 2 | 0.729 | 3.48e-05 | 20.5 | 240 | 12000 | 40 | PF | REFORMATE | 240 | 20.5 | 2/27/2022 12:00:00 AM | 2/27/2016 12:00:00 AM | Active
2 | P00004 | R501 | FLOATING ROOF | NORTH | 39019 | 14293589 | 15375 | 12807 | -2 | 0.7277 | 3.48e-05 | 20.6 | 240 | 12000 | 42 | HT | NAPHTHA INTERMEDIA | 240 | 20.6 | 2/28/2022 12:00:00 AM | 2/29/2016 12:00:00 AM | Active
2 | P00005 | R501 | FLOATING ROOF | NORTH | 33999 | 10862661 | 15343 | 12806 | -13 | 0.75 | 3.48e-05 | 14.9 | 300 | 12000 | 30 | HT | NAPHTHA INTERMEDIA | 300 | 14.9 | 5/17/2022 12:00:00 AM | 5/17/2016 12:00:00 AM | Active
1 | P00007 | R501 | FIXED ROOF | NORTH | 24003 | 5433696 | 13375 | 12410 | 45 | 0.72 | 3.48e-05 | 35 | 400 | 12000 | 35 | HT | NAPHTHA INTERMEDIA | 400 | 35 | 2/23/2024 12:00:00 AM | 2/23/2018 12:00:00 AM | Active
1 | P00008 | R501 | FIXED ROOF | NORTH | 24000 | 5477668 | 13385 | 12379 | 71 | 0.8323 | 3.48e-05 | 0 | 450 | 12100 | 40 | HT | NAPHTHA INTERMEDIA | 450 | 0 | 2/26/2022 12:00:00 AM | 2/26/2016 12:00:00 AM | Active
1 | P00031 | R501 | FIXED ROOF | NORTH | 32007 | 9553633 | 14734 | 13712 | 54 | 0.8047 | 3.48e-05 | 0 | 0 | 13500 | 44 | AV | AVTUR | 450 | 0 | 2/28/2019 12:00:00 AM | 2/20/2013 12:00:00 AM | Active
1 | P00032 | R501 | FIXED ROOF | NORTH | 31993 | 11094240 | 15095 | 14071 | -34 | 0.8119 | 3.48e-05 | 36 | 450 | 13800 | 36 | AV | AVTUR | 450 | 36 | 9/15/2020 12:00:00 AM | 9/15/2014 12:00:00 AM | Active
1 | R00001 | R501 | FIXED ROOF | NORTH | 35984 | 12284128 | 12773 | 12175 | 67 | 0.82 | 3.48e-05 | 0 | 0 | 12000 | 60 | M3 | CRUDE OIL,MUDI | 300 | 0 | 7/31/2019 12:00:00 AM | 7/31/2013 12:00:00 AM | Active
1 | R00002 | R501 | FIXED ROOF | NORTH | 35980 | 11986040 | 12605 | 12107 | 161 | 0.74 | 3.48e-05 | 60 | 300 | 11850 | 60 | WR | CRUDE OIL,WARUKIN TAPIAN TIMUR | 300 | 60 | 11/13/2023 12:00:00 AM | 12/13/2017 12:00:00 AM | Active
1 | R00003 | R501 | FIXED ROOF | NORTH | 35950 | 12088356 | 12692 | 12215 | 123 | 0.7437 | 3.48e-05 | 0 | 850 | 11900 | 33 | M3 | CRUDE OIL,MUDI | 850 | 0 | 2021-02-05 | 2015-02-05 | Active
1 | R00004 | R501 | FIXED ROOF | NORTH | 36027 | 12343890 | 13438 | 12443 | -35 | 0.8735 | 3.48e-05 | 0 | 700 | 12100 | 80 | M3 | CRUDE OIL,MUDI | 700 | 0 | 11/17/2021 12:00:00 AM | 12/17/2015 12:00:00 AM | Active
1 | R00005 | R501 | FIXED ROOF | NORTH | 36008 | 12230825 | 13258 | 12441 | 23 | 0.8735 | 3.48e-05 | 0 | 850 | 12000 | 55 | AU | ARJUNA CRUDE | 850 | 0 | 2021-07-05 | 2015-07-05 | Active
1 | R00006 | R501 | FIXED ROOF | NORTH | 36015 | 12395834 | 13058 | 12341 | 118 | 0.87 | 3.48e-05 | 0 | 320 | 12100 | 45 | AU | ARJUNA CRUDE | 320 | 0 | 5/13/2022 12:00:00 AM | 5/13/2016 12:00:00 AM | Active
1 | R00007 | R501 | FIXED ROOF | NORTH | 36002 | 10137171 | 11545 | 10602 | 55 | 0.8604 | 3.48e-05 | 100 | 200 | 9950 | 100 | HS | SLOP | 200 | 100 | 6/14/2020 12:00:00 AM | 6/14/2014 12:00:00 AM | Active
1 | R00008 | R501 | FIXED ROOF | NORTH | 36022 | 12102156 | 13122 | 12243 | 53 | 0.8514 | 3.48e-05 | 0 | 0 | 11800 | 60 | BU | CRUDE OIL,BUNYU | 800 | 0 | 9/30/2019 12:00:00 AM | 9/21/2013 12:00:00 AM | Active
1 | R00009 | R501 | FIXED ROOF | NORTH | 36006 | 12494274 | 13362 | 12443 | 78 | 0.8735 | 3.48e-05 | 0 | 400 | 12200 | 55 | BU | CRUDE OIL,BUNYU | 400 | 0 | 11/18/2021 12:00:00 AM | 12/18/2015 12:00:00 AM | Active
1 | R00010 | R501 | FIXED ROOF | NORTH | 36559 | 10903561 | 11340 | 10795 | 30 | 0.8735 | 3.48e-05 | 0 | 0 | 10400 | 55 | HS | SLOP | 700 | 0 | 3/28/2021 12:00:00 AM | 3/28/2015 12:00:00 AM | Active
1 | R00011 | R501 | FIXED ROOF | NORTH | 44760 | 17936501 | 12324 | 11785 | 65 | 0.8674 | 3.48e-05 | 0 | 0 | 11500 | 55 | TC | CRUDE OIL,TANJUNG / KAMBITIN | 290 | 0 | 1/31/2020 12:00:00 AM | 1/22/2014 12:00:00 AM | Active
1 | T00001 | R501 | FIXED ROOF | DISWAX | 8533 | 274529 | 5087 | 4969 | 22 | 0.8438 | 3.48e-05 | 0 | 290 | 4800 | 46 | EM | EMPTY TANK | 290 | 0 | 11/16/2021 12:00:00 AM | 12/16/2015 12:00:00 AM | Non Active
1 | T00002 | R501 | FIXED ROOF | DISWAX | 8553 | 275846 | 5113 | 4970 | 6 | 0.86 | 3.48e-05 | 0 | 0 | 4800 | 30 | EM | EMPTY TANK | 0 | 0 | 1/31/2014 12:00:00 AM | 1/27/2008 12:00:00 AM | Non Active
1 | T00201 | R501 | FIXED ROOF | DISWAX | 13707 | 1842793 | 13570 | 12810 | -26 | 0.795 | 3.48e-05 | 70 | 750 | 12500 | 70 | SQ | LAWS 5,BULK | 750 | 70 | 6/15/2020 12:00:00 AM | 6/15/2014 12:00:00 AM | Active
1 | T00202 | R501 | FIXED ROOF | DISWAX | 10666 | 651587 | 7672 | 7380 | 8 | 0.8188 | 3.48e-05 | 0 | 0 | 7300 | 33 | SQ | LAWS 5,BULK | 450 | 0 | 1/31/2020 12:00:00 AM | 1/19/2014 12:00:00 AM | Active
1 | T00203 | R501 | FIXED ROOF | DISWAX | 10672 | 652494 | 7689 | 7380 | 216 | 0.8188 | 3.48e-05 | 0 | 0 | 7300 | 60 | SQ | LAWS 5,BULK | 450 | 0 | 1/31/2020 12:00:00 AM | 1/19/2014 12:00:00 AM | Active
2 | T00900 | R501 | FLOATING ROOF | SOUTH | 36610 | 14399724 | 15100 | 14665 | 40 | 0.8487 | 3.48e-05 | 32 | 400 | 13700 | 32 | TH | TANGGUH CONDENSATE | 400 | 32 | 2023-09-11 | 2017-09-12 | Active
1 | T00901 | R501 | FIXED ROOF | SOUTH | 34743 | 11175571 | 12645 | 12100 | 53 | 0.8557 | 3.48e-05 | 0 | 280 | 11800 | 40 | AU | ARJUNA CRUDE | 280 | 0 | 2/25/2022 12:00:00 AM | 2/25/2016 12:00:00 AM | Active
1 | T00902 | R501 | FIXED ROOF | SOUTH | 34850 | 10820312 | 12678 | 11820 | 80 | 0.863 | 3.48e-05 | 0 | 0 | 11500 | 36 | AU | ARJUNA CRUDE | 500 | 0 | 1/31/2020 12:00:00 AM | 1/21/2014 12:00:00 AM | Active
1 | T00903 | R501 | FIXED ROOF | SOUTH | 23487 | 4940731 | 12184 | 11609 | 103 | 0.8787 | 3.48e-05 | 32 | 750 | 11400 | 32 | S2 | CRUDE OIL,SANGA-SANGA / SAMBOJA | 750 | 32 | 2023-08-11 | 2017-08-12 | Active
1 | T00904 | R501 | FIXED ROOF | SOUTH | 34769 | 8157185 | 9528 | 8936 | 97 | 0.8488 | 3.48e-05 | 0 | 360 | 8600 | 35 | S2 | CRUDE OIL,SANGA-SANGA / SAMBOJA | 360 | 0 | 2021-04-05 | 2015-04-05 | Active
1 | U00010 | R501 | FIXED ROOF | UTILITIES | 8992 | 552553 | 9534 | 9099 | 237 | 0.8735 | 3.48e-05 | 60 | 540 | 8700 | 55 | AD | ADO/SOLAR-48 | 540 | 60 | 2021-11-10 | 2015-11-11 | Active
1 | U00011 | R501 | FIXED ROOF | UTILITIES | 8993 | 560971 | 9502 | 9085 | 215 | 0.9 | 3.48e-05 | 0 | 680 | 8800 | 55 | AD | ADO/SOLAR-48 | 680 | 0 | 11/16/2021 12:00:00 AM | 12/16/2015 12:00:00 AM | Active
1 | U00012 | R501 | FIXED ROOF | UTILITIES | 9007 | 560187 | 9335 | 9126 | 3 | 0.792 | 3.48e-05 | 47 | 620 | 8800 | 47 | AF | ADO REF.FUEL | 620 | 47 | 2020-12-11 | 2014-12-11 | Active
1 | U00013 | R501 | FIXED ROOF | UTILITIES | 24022 | 5697193 | 14080 | 12797 | -20 | 0.9 | 3.48e-05 | 0 | 0 | 12500 | 55 | OR | REFINERY FUEL OIL | 0 | 0 | 9/30/2019 12:00:00 AM | 9/26/2013 12:00:00 AM | Active
1 | V00001 | R501 | FIXED ROOF | SOUTH | 35951 | 4956107 | 5845 | 5362 | 50 | 0.868 | 3.48e-05 | 0 | 0 | 5000 | 65 | HS | SLOP | 400 | 0 | 3/30/2021 12:00:00 AM | 3/30/2015 12:00:00 AM | Active
1 | V00002 | R501 | FIXED ROOF | SOUTH | 11601 | 896813 | 9200 | 8758 | 42 | 0.8093 | 3.48e-05 | 0 | 0 | 8500 | 34 | HS | SLOP | 250 | 0 | 1/31/2019 12:00:00 AM | 1/16/2013 12:00:00 AM | Active
1 | V00003 | R501 | FIXED ROOF | NORTH | 15985 | 2416572 | 12699 | 12320 | 83 | 0.8707 | 3.48e-05 | 0 | 0 | 12000 | 43 | HS | SLOP | 600 | 0 | 1/31/2019 12:00:00 AM | 2013-08-01 | Active
1 | V00004 | R501 | FIXED ROOF | NORTH | 16004 | 2450406 | 12625 | 12285 | 80 | 0.8663 | 3.48e-05 | 0 | 320 | 12100 | 50 | HS | SLOP | 320 | 0 | 2023-07-11 | 2017-07-12 | Active
1 | V00005 | R501 | FIXED ROOF | NORTH | 7497 | 312295 | 7937 | 7415 | 16 | 0.8734 | 3.48e-05 | 35 | 400 | 7100 | 35 | EM | EMPTY TANK | 400 | 35 | 11/13/2020 12:00:00 AM | 11/13/2014 12:00:00 AM | Active
1 | V00006 | R501 | FIXED ROOF | NORTH | 7499 | 313765 | 7925 | 7408 | 50 | 0.84 | 3.48e-05 | 30 | 400 | 7100 | 30 | HS | SLOP | 400 | 30 | 5/15/2022 12:00:00 AM | 5/15/2016 12:00:00 AM | Active
1 | V00007 | R501 | FIXED ROOF | NORTH | 7488 | 308320 | 7759 | 7267 | -30 | 0.8093 | 3.48e-05 | 0 | 0 | 7000 | 40 | SA | CRUDE OIL,SANGATTA | 290 | 0 | 1/31/2019 12:00:00 AM | 1/17/2013 12:00:00 AM | Active
1 | V00010 | R501 | FIXED ROOF | DISWAX | 7461 | 306131 | 7389 | 7228 | 5 | 0.8734 | 3.48e-05 | 35 | 300 | 7000 | 35 | HS | SLOP | 300 | 35 | 11/13/2020 12:00:00 AM | 11/13/2014 12:00:00 AM | Active
1 | V00011 | R501 | FIXED ROOF | DISWAX | 11993 | 573980 | 5743 | 5342 | 39 | 0.8734 | 3.48e-05 | 35 | 360 | 5100 | 35 | HS | SLOP | 360 | 35 | 11/13/2020 12:00:00 AM | 11/13/2014 12:00:00 AM | Active
1 | V00013 | R501 | FIXED ROOF | DISWAX | 3227 | 30400 | 4043 | 3977 | -6 | 0.819 | 3.48e-05 | 75 | 300 | 3700 | 75 | HS | SLOP | 300 | 75 | 6/16/2020 12:00:00 AM | 6/16/2014 12:00:00 AM | Active
1 | V00016 | R501 | FIXED ROOF | DISWAX | 4577 | 84432 | 5512 | 5392 | -7 | 0.9181 | 3.48e-05 | 0 | 0 | 5120 | 70 | HS | SLOP | 0 | 0 | 9/30/2019 12:00:00 AM | 9/26/2013 12:00:00 AM | Active
1 | V00017 | R501 | FIXED ROOF | DISWAX | 4497 | 91281 | 6295 | 5950 | 0 | 0 | 3.48e-05 | 0 | 0 | 5750 | 50 | HS | SLOP | 0 | 0 | 11/30/1998 12:00:00 AM |  | Active
1 | V0017B | R501 | FIXED ROOF | DISWAX | 6097 | 207613 | 8182 | 7310 | 108 | 0.925 | 3.48e-05 | 0 | 0 | 7100 | 70 | HS | SLOP | 200 | 0 | 1/31/2020 12:00:00 AM | 1/18/2014 12:00:00 AM | Active
1 | WRHOUSE | R501 | FIXED ROOF |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | RJ | YELLOW BATIK (26BYS) | 0 | 0 | 12/31/2099 12:00:00 AM |  | Active
1 | YARD | R501 | FIXED ROOF |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | CK |  | 0 | 0 | 12/31/2099 12:00:00 AM |  | Active

## Heαder Tαnk (hidden)
HEADER TANK
Petunjuk Pengisian Formulir:
| 1 | Isi field hanya pada data yang ingin diubah saja.
| 2 | Jika ingin mengosongkan data pada web ROAS maka kosongkan field dengan manghapus kata 'TIDAK BERUBAH'.
| Plant |  | R | 5 | 0 | 1 | [7 empty] | Product | TIDAK BERUBAH
| Tank |  | T | 0 | 0 | 9 | 0 | 2 | [5 empty] | Product Code | TIDAK BERUBAH
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
| Jenis Atap |  | 1: T; 2: e; 3: t; 4: a; 5: p | Tera Expired | O | k | t | o | b | e | r | 2 | 0 | 2 | 9
| Area |  | 1: S; 2: O; 3: U; 4: T; 5: H | Last Tera Date | 4 | 3 | 7 | 4 | 9
| Unit |  | B | A | L | I | K | P | A | P | A | N |  | Calibration Status | TIDAK BERUBAH
| Status Tank |  | B: ü; A: Active; P: Not Active |  | Last Calibrate By | TIDAK BERUBAH
| Diameter |  | B: 3; A: 4; L: 8; I: 2; K: 0 |  | Faktor Koreksi Volume | 3 | 6
| Tank Height |  | B: 1; A: 1; L: 8; I: 1; K: 8 |  | Valve Factor | TIDAK BERUBAH
| Total Tank Height |  | B: 1; A: 1; L: 8; I: 1; K: 8 |  | Paint Factor | TIDAK BERUBAH
| Datum Dip |  | B: 8; A: 2 |  | Clinggage Factor | TIDAK BERUBAH
| Dip Hole |  | B: 1; A: 2; L: 6; I: 7; K: 1 |  | Koefisien Muai | TIDAK BERUBAH
| Drawable Height |  | B: 5; A: 0; L: 0 |  | Koefisien Jenis Minyak | TIDAK BERUBAH
| Massa Jenis |  | B: 0; A: .; L: 8; I: 6 |  | Koefisien Type Atap | TIDAK BERUBAH
| Suhu |  | B: 3; A: 6 |  | Tag PI Pressure | TIDAK BERUBAH
| Volume Bersih |  | 1 | 0 | 7 | 8 | 3 | 2 | 7 | 9 |  |  |  | Tag PI Temperature | TIDAK BERUBAH
| Min. Tinggi Cairan |  | B: 5; A: 0; L: 0 |  | Tag PI Level | TIDAK BERUBAH
| Max Tinggi Cairan |  | B: 1; A: 1; L: 5; I: 0; K: 0
[21 empty] | Diusulkan Oleh : |  |  |  |  | Diterima Oleh :
[21 empty] | Legal Requestor |  |  |  |  | Unit MDM
[21 empty] | Tgl : |  |  |  |  | Tgl :
Arief Maukar R.

## Fraction Table (hidden)
FRACTION TABLE
| Plant |  | R501
| Tank |  | T00902
| Tank Height |  | 11818
| Plant | Tank | Dari (cm) | Sampai (cm) | Tinggi (mm) | Volume
| R501 | T00902 | 50 | 168 | 1 | 951.3 | Diusulkan Oleh : | Diterima Oleh :
| R501 | T00902 | 50 | 168 | 2 | 1902.6 | Legal Requestor | Unit MDM
| R501 | T00902 | 50 | 168 | 3 | 2854 | Tgl : | Tgl :
| R501 | T00902 | 50 | 168 | 4 | 3805.3
| R501 | T00902 | 50 | 168 | 5 | 4756.6
| R501 | T00902 | 50 | 168 | 6 | 5707.9 | Arief Maukar R.
| R501 | T00902 | 50 | 168 | 7 | 6659.2
| R501 | T00902 | 50 | 168 | 8 | 7610.6
| R501 | T00902 | 50 | 168 | 9 | 8561.9
| R501 | T00902 | 50 | 168 | 10 | 9513.2
| R501 | T00902 | 168 | 291 | 1 | 949.3
| R501 | T00902 | 168 | 291 | 2 | 1898.6
| R501 | T00902 | 168 | 291 | 3 | 2847.9
| R501 | T00902 | 168 | 291 | 4 | 3797.2
| R501 | T00902 | 168 | 291 | 5 | 4746.5
| R501 | T00902 | 168 | 291 | 6 | 5695.8
| R501 | T00902 | 168 | 291 | 7 | 6645.1
| R501 | T00902 | 168 | 291 | 8 | 7594.4
| R501 | T00902 | 168 | 291 | 9 | 8543.7
| R501 | T00902 | 168 | 291 | 10 | 9493
| R501 | T00902 | 291 | 445 | 1 | 951.5
| R501 | T00902 | 291 | 445 | 2 | 1903.1
| R501 | T00902 | 291 | 445 | 3 | 2854.6
| R501 | T00902 | 291 | 445 | 4 | 3806.2
| R501 | T00902 | 291 | 445 | 5 | 4757.7
| R501 | T00902 | 291 | 445 | 6 | 5709.3
| R501 | T00902 | 291 | 445 | 7 | 6660.8
| R501 | T00902 | 291 | 445 | 8 | 7612.4
| R501 | T00902 | 291 | 445 | 9 | 8563.9
| R501 | T00902 | 291 | 445 | 10 | 9515.5
| R501 | T00902 | 445 | 566 | 1 | 949.5
| R501 | T00902 | 445 | 566 | 2 | 1899.1
| R501 | T00902 | 445 | 566 | 3 | 2848.6
| R501 | T00902 | 445 | 566 | 4 | 3798.1
| R501 | T00902 | 445 | 566 | 5 | 4747.7
| R501 | T00902 | 445 | 566 | 6 | 5697.2
| R501 | T00902 | 445 | 566 | 7 | 6646.7
| R501 | T00902 | 445 | 566 | 8 | 7596.3
| R501 | T00902 | 445 | 566 | 9 | 8545.8
| R501 | T00902 | 445 | 566 | 10 | 9495.3
| R501 | T00902 | 566 | 751 | 1 | 950.4
| R501 | T00902 | 566 | 751 | 2 | 1900.7
| R501 | T00902 | 566 | 751 | 3 | 2851.1
| R501 | T00902 | 566 | 751 | 4 | 3801.4
| R501 | T00902 | 566 | 751 | 5 | 4751.8
| R501 | T00902 | 566 | 751 | 6 | 5702.1
| R501 | T00902 | 566 | 751 | 7 | 6652.5
| R501 | T00902 | 566 | 751 | 8 | 7602.9
| R501 | T00902 | 566 | 751 | 9 | 8553.2
| R501 | T00902 | 566 | 751 | 10 | 9503.6
| R501 | T00902 | 751 | 934 | 1 | 950.8
| R501 | T00902 | 751 | 934 | 2 | 1901.7
| R501 | T00902 | 751 | 934 | 3 | 2852.5
| R501 | T00902 | 751 | 934 | 4 | 3803.3
| R501 | T00902 | 751 | 934 | 5 | 4754.1
| R501 | T00902 | 751 | 934 | 6 | 5705
| R501 | T00902 | 751 | 934 | 7 | 6655.8
| R501 | T00902 | 751 | 934 | 8 | 7606.6
| R501 | T00902 | 751 | 934 | 9 | 8557.4
| R501 | T00902 | 751 | 934 | 10 | 9508.3
| R501 | T00902 | 934 | 1118 | 1 | 952
| R501 | T00902 | 934 | 1118 | 2 | 1904
| R501 | T00902 | 934 | 1118 | 3 | 2856
| R501 | T00902 | 934 | 1118 | 4 | 3808
| R501 | T00902 | 934 | 1118 | 5 | 4759.9
| R501 | T00902 | 934 | 1118 | 6 | 5711.9
| R501 | T00902 | 934 | 1118 | 7 | 6663.9
| R501 | T00902 | 934 | 1118 | 8 | 7615.9
| R501 | T00902 | 934 | 1118 | 9 | 8567.9
| R501 | T00902 | 934 | 1118 | 10 | 9519.9
| R501 | T00902 | 1118 | 1150 | 1 | 952.7
| R501 | T00902 | 1118 | 1150 | 2 | 1903.3
| R501 | T00902 | 1118 | 1150 | 3 | 2858
| R501 | T00902 | 1118 | 1150 | 4 | 3810.6
| R501 | T00902 | 1118 | 1150 | 5 | 4763.3
| R501 | T00902 | 1118 | 1150 | 6 | 5715.9
| R501 | T00902 | 1118 | 1150 | 7 | 6668.6
| R501 | T00902 | 1118 | 1150 | 8 | 7621.2
| R501 | T00902 | 1118 | 1150 | 9 | 8573.9
| R501 | T00902 | 1118 | 1150 | 10 | 9526.5